## Request a quote
Request a quotation for 301-KAP50 / 301-KAP50-S
Specify 301-KAP50 (polyimide outer) or 301-KAP50-S (316L SS outer braid), coil length, or pre-terminated UHV PEEK-insulated SMA / SMB / Microdot assemblies.
Enter the length you need next to each size in the Sizes sheet (column A) and send this file to:
Email: info@tebaoflex.com
WhatsApp / phone: +86 185 9591 3361
Online quote form: https://tebaoflex.com/cables/301-kap50-301-kap50-s-polyimide-uhv-50-ohm-coaxial-cable/#template-selection
An application engineer replies within 2 business hours.

## Sizes
301-KAP50 / 301-KAP50-S Polyimide UHV 50 Ohm Coaxial Cable – Dimensions & construction (4)
TEBAOFLEX ¦ https://tebaoflex.com/cables/301-kap50-301-kap50-s-polyimide-uhv-50-ohm-coaxial-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on UHV Coaxial Spec (≤ 10⁻¹⁰ mbar) / MIL-DTL-17 Ref.. Values for your order are confirmed in the quotation.
Enter the length you need next to each size and send this file to info@tebaoflex.com or WhatsApp +86 185 9591 3361.
Your length (m) | Order designation | Construction / Variant | Conductor | Outer Ø ≈ mm | Metal kg/km | Weight ≈ kg/km | Min. bend R fixed mm | Series
| 301-KAP50 Smooth Polyimide Outer (Ø 1.95 mm) | 50 Ω Ag-Cu Braid + Polyimide Jacket | 24 AWG (19×0.12 Ag-Cu) | 1.95 | 6.2 | 10.5 | 10 | Standard 301-KAP50 / 301-KAP50-S variants
| 301-KAP50-S With 316L SS Outer Braid (Ø 2.25 mm) | 50 Ω Ag-Cu Braid + Polyimide + 316L SS | 24 AWG (19×0.12 Ag-Cu) | 2.25 | 6.2 | 16.8 | 11 | Standard 301-KAP50 / 301-KAP50-S variants
| 310-PEEK50-TRIAX PEEK UHV Triaxial Sister | 50 Ω Dual Guard Braids (PEEK UHV) | 26 AWG (Ø 0.48 Ag-Cu) | 2.95 | 11.5 | 19.5 | 15 | Standard 301-KAP50 / 301-KAP50-S variants
| RG-196A/U Non-Radiation PTFE Coax (Ø 2.03 mm) | 50 Ω Single SPC Braid (PTFE) | 30 AWG (7×0.10 SCCS) | 2.03 | 4.2 | 8.9 | 10 | Standard 301-KAP50 / 301-KAP50-S variants
Dimensions and weights are nominal production values subject to manufacturing tolerances per UHV Coaxial Spec (≤ 10⁻¹⁰ mbar) / MIL-DTL-17 Ref..

## Attenuation & RF power handling
301-KAP50 / 301-KAP50-S Polyimide UHV 50 Ohm Coaxial Cable – Attenuation & RF power handling (6)
TEBAOFLEX ¦ https://tebaoflex.com/cables/301-kap50-301-kap50-s-polyimide-uhv-50-ohm-coaxial-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on UHV Coaxial Spec (≤ 10⁻¹⁰ mbar) / MIL-DTL-17 Ref.. Values for your order are confirmed in the quotation.
Frequency | Nominal attenuation dB / 100 m | Nominal attenuation dB / 100 ft | Max. CW RF power W (@ 40 °C) | Impedance Z₀ Ω | Velocity VoP %
10 MHz | 6.5 | 1.98 | 180 | 50 | 65
100 MHz | 21 | 6.4 | 60 | 50 | 65
500 MHz | 49 | 14.94 | 26 | 50 | 65
1000 MHz (1 GHz) | 72 | 21.95 | 18 | 50 | 65
3000 MHz (3 GHz) | 135 | 41.15 | 10 | 50 | 65
5000 MHz (5 GHz) | 185 | 56.39 | 7 | 50 | 65
Attenuation values apply at +20 °C ambient into a matched 50 Ω load; CW power ratings assume VSWR 1.0:1 at +40 °C ambient.

## Technical data
301-KAP50 / 301-KAP50-S Polyimide UHV 50 Ohm Coaxial Cable – Technical data
TEBAOFLEX ¦ https://tebaoflex.com/cables/301-kap50-301-kap50-s-polyimide-uhv-50-ohm-coaxial-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on UHV Coaxial Spec (≤ 10⁻¹⁰ mbar) / MIL-DTL-17 Ref.. Values for your order are confirmed in the quotation.
Group | Parameter | Value
Key data | Vacuum rating | ≤ 10⁻¹⁰ mbar UHV (TML < 0.3 % · CVCM < 0.02 %)
Key data | Impedance Z₀ | 50 Ω ± 2.5 Ω (DC to 5 GHz (Ø 1.95 / 2.25 mm))
Key data | Temp. range | -269 to +250 °C (4 K cryogenic to +250 °C bakeout)
Key data | Radiation dose | Up to 10⁷ Gy (Halogen-free aromatic polyimide)
Key data | Velocity / Cap. | 65 % · 105 pF/m (0.60 mm (19×0.12) Ag-Cu center)
Key data | Outer shield (-S) | AISI 316L SS (Abrasion & grounding guard)
Electrical | Characteristic impedance Z₀ | 50 Ω ± 2.5 Ω
Electrical | Nominal capacitance | 105.0 pF/m (32.0 pF/ft)
Electrical | Velocity of propagation (VoP) | 65.0 % (5.13 ns/m delay)
Electrical | Center / shield DC resistance | ≤ 86 Ω/km / ≤ 26 Ω/km at 20 °C
Electrical | Max. operating voltage | 1,200 V DC / 850 V rms AC (in air/UHV)
Electrical | Dielectric withstand test | 2.0 kV rms at 50/60 Hz
Thermal | Operating temperature range | -269 °C (4 K) to +250 °C continuous
Thermal | Short-term bakeout peak | +300 °C in vacuum
Thermal | Cryogenic thermal shock | Direct immersion in LN₂ (77 K) & LHe (4 K)
Thermal | Outgassing (ASTM E595 / ECSS) | TML < 0.30 % · CVCM < 0.01 % (post-bake)
Thermal | UHV pressure compatibility | ≤ 1 × 10⁻¹⁰ mbar (10⁻⁸ Pa)
Thermal | Flammability in air | Self-extinguishing UL 94 V-0
Mechanical | Center conductor stranding | 19 × 0.12 mm Ag-OFC (Ø 0.60 mm, 24 AWG)
Mechanical | Dielectric diameter | 1.48 mm ± 0.05 mm
Mechanical | Outer diameter (KAP50 / KAP50-S) | 1.95 mm ± 0.08 mm / 2.25 mm ± 0.10 mm
Mechanical | Min. bend radius (fixed / dynamic) | 10 mm (5×D) / 20 mm (10×D)
Mechanical | Nominal weight (KAP50 / KAP50-S) | 10.5 kg/km / 16.8 kg/km
Mechanical | Connector compatibility | UHV peek/ceramic SMA, SMB, Microdot, BNC
Resistance | Ionizing radiation resistance | Up to 10⁷ Gy (10 MGy / 10⁹ rad) gamma/e⁻
Resistance | Halogen & fluorine free | Zero F⁻ outgassing under beam bombardment
Resistance | Magnetic permeability | Non-magnetic Ag-Cu & annealed 316L SS
Resistance | Solvent & ultrasonic cleaning | Compatible with acetone, IPA & UHV clean
Resistance | Abrasion protection (-S version) | 316L stainless braid resists sharp ports
Resistance | RoHS / REACH status | RoHS 3 compliant
Standards | Vacuum outgassing standard | ASTM E595 / NASA SP-R-0022A / ECSS-Q-ST-70-02
Standards | Radiation endurance reference | CERN / DESY / ESRF beamline qualification
Standards | Coaxial RF standard | IEC 61196-1
Standards | Conductor standard | ASTM B298 oxygen-free silver-plated Cu
Standards | Stainless overbraid standard | EN 10088 / AISI 316L non-magnetic
Standards | Quality certification | UHV cleanroom certificate + EN 10204 3.1
Identification | Dielectric colour | Amber-gold translucent polyimide
Identification | Outer appearance (301-KAP50) | Amber-gold polyimide film
Identification | Outer appearance (301-KAP50-S) | Bright metallic AISI 316L stainless braid
Identification | Packaging | UHV-cleaned, double-bagged in cleanroom PE
Identification | Standard lengths | 5 m, 10 m, 25 m, 50 m, 100 m
Identification | Batch traceability | Cleanroom lot number & HV/TDR test sheet

## Construction
301-KAP50 / 301-KAP50-S Polyimide UHV 50 Ohm Coaxial Cable – Construction
TEBAOFLEX ¦ https://tebaoflex.com/cables/301-kap50-301-kap50-s-polyimide-uhv-50-ohm-coaxial-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on UHV Coaxial Spec (≤ 10⁻¹⁰ mbar) / MIL-DTL-17 Ref.. Values for your order are confirmed in the quotation.
No. | Layer | Material | Reference
1 | Stranded Ag-OFC center conductor (24 | 19 × 0.12 mm (Ø 0.60 mm, 0.22 mm²) ultra-clean silver-plated oxygen-free copper
2 | Multi-layer fused polyimide dielectric | Precision wrapped and thermally sealed aromatic polyimide film, Ø 1.48 mm (εr ≈ 3.4, 65
3 | Silver-plated copper RF shield braid | 0.10 mm silver-plated oxygen-free copper wire braid (≥ 90 % coverage)
4 | Fused amber polyimide UHV jacket | Sealed aromatic polyimide outer sheath, amber-gold (Ø 1.95 mm on 301-KAP50)
5 | AISI 316L stainless-steel overbraid | Non-magnetic AISI 316L stainless-steel wire outer braid (Ø 2.25 mm on 301-KAP50-S) for
| 301-KAP50 (Smooth Polyimide Outer Jacket, Ø 1.95 mm) | Electrically isolated amber polyimide outer surface so the RF shield floats from vacuum chamber walls.
| 301-KAP50-S (316L Stainless-Steel Overbraided, Ø 2.25 mm) | Adds an external non-magnetic 316L stainless-steel braid for flange chafe protection and static charge bleed.

## Quality and supply
301-KAP50 / 301-KAP50-S Polyimide UHV 50 Ohm Coaxial Cable – Testing, approvals and supply
TEBAOFLEX ¦ https://tebaoflex.com/cables/301-kap50-301-kap50-s-polyimide-uhv-50-ohm-coaxial-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on UHV Coaxial Spec (≤ 10⁻¹⁰ mbar) / MIL-DTL-17 Ref.. Values for your order are confirmed in the quotation.
Item | Detail | Frequency / scope
Routine test | TDR characteristic impedance & 2.0 kV HV test — 50 Ω ± 2.5 Ω; 2.0 kV rms AC without breakdown | Every coil
Routine test | Hydrocarbon-free UHV ultrasonic cleaning check — Certified free of drawing lubricants, silicones, and halogens | Every coil
Sample test | Residual gas analysis (RGA) & ASTM E595 outgassing — Compatible with ≤ 10⁻¹⁰ mbar UHV after +200 °C bakeout | Per batch
Type test | Liquid helium (4 K) to +250 °C thermal shock — Zero polyimide delamination or shield fracture over 50 cycles | Qualification
Standard / approval | UHV ≤ 10⁻¹⁰ mbar Certified
Standard / approval | ASTM E595 / ECSS-Q-ST-70-02
Standard / approval | 10 MGy Radiation-Hard
Company certificate | VDE
Company certificate | UL
Company certificate | cUL
Company certificate | CE
Company certificate | KTL
Company certificate | MSHA
Company certificate | MA
Company certificate | ISO 9001
Company certificate | ISO 14001
Company certificate | ISO 45001
Sheath marking | TEBAOFLEX 301-KAP50 / 301-KAP50-S 50 OHM
Supply | Delivery form: Spools, reels or export wooden drums
Supply | Standard reel lengths: 100 m · 305 m (1000 ft) · 500 m · 1000 m
Supply | MOQ: 100 m (stock) · 305–500 m (production)
Supply | Lead time: 1–3 weeks ex-works
Supply | Incoterms: EXW, FCA, FOB, CIF, DAP