## Request a quote
Request a quotation for HF SABmed S
Medical & electrosurgical cable engineer
Enter the length you need next to each size in the Sizes sheet (column A) and send this file to:
Email: info@tebaoflex.com
WhatsApp / phone: +86 185 9591 3361
Online quote form: https://tebaoflex.com/cables/hf-sabmed-s-autoclavable-silicone-electrosurgical-high-frequency-medical-cable/#template-selection
An application engineer replies within 2 business hours.

## Sizes
HF SABmed S Autoclavable Silicone Electrosurgical High-Frequency – Configuration, cross-section, HF kVp rating & OD (12)
TEBAOFLEX ¦ https://tebaoflex.com/cables/hf-sabmed-s-autoclavable-silicone-electrosurgical-high-frequency-medical-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 60601-2-2 / ISO 10993-1 / USP Class VI / DIN EN ISO 17665-1. Values for your order are confirmed in the quotation.
Enter the length you need next to each size and send this file to info@tebaoflex.com or WhatsApp +86 185 9591 3361.
Your length (m) | Order designation | Cores × Cross-Section | Stranding (Ø 0.05 mm) | Max HF U_p kVp | Test U kV AC | Nom. OD mm | Weight kg/km | Series
| HF-SABMED-S-1X0.75-5KV | 1 × 0.75 mm² Monopolar | 384 × 0.05 mm SnCu | 5 | 6 | 3.6 | 19.5 | Monopolar Double-Insulated & Bipolar Figure-8 / Round Electrosurgical Cables
| HF-SABMED-S-1X1.00-5KV | 1 × 1.00 mm² Monopolar | 512 × 0.05 mm SnCu | 5 | 6 | 3.9 | 23.8 | Monopolar Double-Insulated & Bipolar Figure-8 / Round Electrosurgical Cables
| HF-SABMED-S-1X1.50-5KV | 1 × 1.50 mm² Monopolar | 768 × 0.05 mm SnCu | 5 | 6 | 4.3 | 30.5 | Monopolar Double-Insulated & Bipolar Figure-8 / Round Electrosurgical Cables
| HF-SABMED-S-2X0.50-TWIN | 2 × 0.50 mm² Fig-8 Twin | 256 × 0.05 mm SnCu | 2.5 | 4 | 2.4 × 4.8 | 21 | Monopolar Double-Insulated & Bipolar Figure-8 / Round Electrosurgical Cables
| HF-SABMED-S-2X0.50-RND | 2 × 0.50 mm² Bipolar Rnd | 256 × 0.05 mm SnCu | 3.5 | 4.5 | 4.6 | 28.5 | Monopolar Double-Insulated & Bipolar Figure-8 / Round Electrosurgical Cables
| HF-SABMED-S-2X0.75-RND | 2 × 0.75 mm² Bipolar Rnd | 384 × 0.05 mm SnCu | 4.5 | 5 | 5.1 | 36.2 | Monopolar Double-Insulated & Bipolar Figure-8 / Round Electrosurgical Cables
| HF-SABMED-S-2X1.00-RND | 2 × 1.00 mm² Bipolar HV | 512 × 0.05 mm SnCu | 5 | 6 | 5.6 | 44 | Monopolar Double-Insulated & Bipolar Figure-8 / Round Electrosurgical Cables
| HF-SABMED-S-3X0.50-SW | 3 × 0.50 mm² Hand-Switch | 256 × 0.05 mm SnCu | 4.5 | 5 | 4.9 | 34.5 | 3-Core, 4-Core & 5-Core Hand-Switched Monopolar / Bipolar / Vessel-Sealer Cables
| HF-SABMED-S-3X0.75-SW | 1×0.75 + 2×0.25 mm² Pencil | 384 / 128 × 0.05 mm | 5 | 6 | 4.8 | 32 | 3-Core, 4-Core & 5-Core Hand-Switched Monopolar / Bipolar / Vessel-Sealer Cables
| HF-SABMED-S-4X0.34-BIP | 4 × 0.34 mm² Bipolar Sw. | 175 × 0.05 mm SnCu | 3.5 | 4.5 | 4.8 | 31.8 | 3-Core, 4-Core & 5-Core Hand-Switched Monopolar / Bipolar / Vessel-Sealer Cables
| HF-SABMED-S-2X0.75+2X0.25 | 2×0.75 + 2×0.25 mm² Sealer | 384 / 128 × 0.05 mm | 4.5 | 5 | 5.4 | 41.5 | 3-Core, 4-Core & 5-Core Hand-Switched Monopolar / Bipolar / Vessel-Sealer Cables
| HF-SABMED-S-2X0.75+4X0.14 | 2×0.75 + 4×0.14 mm² Smart | 384 / 72 × 0.05 mm | 4.5 | 5 | 5.7 | 45.2 | 3-Core, 4-Core & 5-Core Hand-Switched Monopolar / Bipolar / Vessel-Sealer Cables
Custom overmould-ready cut leads with silicone strain reliefs, Valleylab/Erbe/Bowa 3-pin or 2-pin plugs, and RFID instrument ID pairs available.

## Technical data
HF SABmed S Autoclavable Silicone Electrosurgical High-Frequency – Technical data
TEBAOFLEX ¦ https://tebaoflex.com/cables/hf-sabmed-s-autoclavable-silicone-electrosurgical-high-frequency-medical-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 60601-2-2 / ISO 10993-1 / USP Class VI / DIN EN ISO 17665-1. Values for your order are confirmed in the quotation.
Group | Parameter | Value
Key data | HF peak voltage | Up to 5.0 kVp (IEC 60601-2-2 coagulation & cut)
Key data | Autoclave rating | +134°C / 2.1 bar (300–500+ steam cycles (ISO 17665))
Key data | Biocompatibility | ISO 10993-1 / -5 (USP Class VI platinum-cured SiR)
Key data | Configurations | 1C, 2C, 3C, 4C (Monopolar, Bipolar & Hand-Switch)
Key data | Temperature | −50 to +180 °C (Peaks to +200 °C sterilization)
Key data | Surface finish | Non-Stick Matte (Low friction, drape-limp & wipe-clean)
Electrical | Max HF rated peak voltage | 2.5 kVp (bipolar standard) · 4.5 kVp to 5.0 kVp (monopolar / HV spray coag)
Electrical | Routine AC test voltage | 6.0 kV AC (monopolar) · 4.0 kV AC core-to-core & core-to-water (bipolar)
Electrical | Operating RF frequency | 300 kHz to 4.0 MHz (IEC 60601-2-2 electrosurgical generators)
Electrical | Core capacitance (bipolar) | ≤ 48 pF/m core-to-core (minimizes RF leakage current)
Thermal | Continuous temperature range | −50 °C to +180 °C (short-term +200 °C)
Thermal | Steam autoclave sterilization | +134 °C to +138 °C at 2.1–2.3 bar for 5–18 min (300–500+ cycles per ISO 17665-1)
Mechanical | Individual strand diameter | 0.05 mm (44 AWG) ultra-flexible tinned copper
Mechanical | Min. bend radius | 4 × OD (drape-limp zero-memory handling)
Resistance | Biocompatibility | ISO 10993-1 / ISO 10993-5 (non-cytotoxic) · USP Class VI
Resistance | Sterilization & cleaning | Steam autoclave, EtO, H2O2 gas plasma (STERRAD) & pH 10.5 washer-disinfectors
Resistance | Halogen & plasticizer free | 100 % platinum-cured silicone (zero phthalates, latex or peroxide bloom)
Standards | Electrosurgical safety | IEC 60601-2-2 (Clause 201.8.8.3.102–104 HF surgical accessories)
Standards | Medical device quality | ISO 10993-1/-5 · USP Class VI · RoHS / REACH
Identification | Standard medical colours | Surgical Blue (RAL 5015), Light Grey (RAL 7035), White, Black, Red, Green, Yellow
Identification | Customs tariff (HS) | 8544.49 / 9018.90

## Construction
HF SABmed S Autoclavable Silicone Electrosurgical High-Frequency – Construction
TEBAOFLEX ¦ https://tebaoflex.com/cables/hf-sabmed-s-autoclavable-silicone-electrosurgical-high-frequency-medical-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 60601-2-2 / ISO 10993-1 / USP Class VI / DIN EN ISO 17665-1. Values for your order are confirmed in the quotation.
No. | Layer | Material | Reference
1 | Ultra-fine Class 6 or Litz tinned copper (Ø 0.05 mm strands) | Hundreds of hair-fine 0.05 mm tinned copper wires bunched or Litz-stranded in 0.25 mm² to 1.50 mm² cross-sections
2 | High-dielectric-strength HF silicone or fluoropolymer (FEP) core wall | Void-free platinum-cured silicone rubber or FEP fluoropolymer primary insulation engineered for high-frequency crest voltages up to 5.0 kVp
3 | Platinum-cured ISO 10993 / USP Class VI medical silicone jacket | Addition-crosslinked (platinum-catalyzed) biocompatible silicone rubber with permanent non-stick matte surface treatment (zero peroxide by-products)
| HF SABmed S Monopolar (1-Core 0.75–1.5 mm², up to 5.0 kVp) | Double-insulated single-core high-voltage RF electrode lead for monopolar endoscopic, arthroscopic and resectoscope loops
| HF SABmed S Bipolar Twin / Round (2-Core 0.50–1.0 mm², 2.5–4.5 kVp) | Figure-8 twin-lead or round jacketed 2-core cable for reusable bipolar forceps and laparoscopic coagulators
| HF SABmed S Hand-Switched (3-Core & 4-Core Hybrid RF + Cut/Coag Switch) | Combines 1 or 2 high-voltage RF power cores with 2 finger-switch control cores for hand-activated electrosurgical pencils and sealers

## Quality and supply
HF SABmed S Autoclavable Silicone Electrosurgical High-Frequency – Testing, approvals and supply
TEBAOFLEX ¦ https://tebaoflex.com/cables/hf-sabmed-s-autoclavable-silicone-electrosurgical-high-frequency-medical-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 60601-2-2 / ISO 10993-1 / USP Class VI / DIN EN ISO 17665-1. Values for your order are confirmed in the quotation.
Item | Detail | Frequency / scope
Routine test | IEC 60601-2-2 high-frequency & 50 Hz wet dielectric withstand test — 100 % spark test + water-immersion AC/HF breakdown and HF capacitive leakage verification | 100 % of manufactured length
Sample test | +134 °C / 2.1 bar steam-autoclave endurance cycling (300–500 cycles) — Verify zero jacket cracking, tackiness or dielectric loss after repeated prevacuum steam sterilization | Qualification & batch audit
Sample test | ISO 10993-5 cytotoxicity & USP Class VI extractables verification — Certified platinum-cured medical elastomer lot traceability | Every silicone compound batch
Standard / approval | IEC 60601-2-2 High-Frequency Surgical Accessory Compliant
Standard / approval | ISO 10993-1 / -5 & USP Class VI Biocompatible Silicone
Company certificate | VDE
Company certificate | UL
Company certificate | cUL
Company certificate | CE
Company certificate | KTL
Company certificate | MSHA
Company certificate | MA
Company certificate | ISO 9001
Company certificate | ISO 14001
Company certificate | ISO 45001
Sheath marking | Smooth Sterilizable Unprinted Surface (or Biocompatible Laser/Silicone Ink Legend)
Supply | Delivery form: Cleanroom-bagged reels (100 m / 250 m) or pre-cut stripping-ready surgical leads
Supply | Compliance pack: ISO 10993-5 CoC, USP Class VI compound cert & IEC 60601-2-2 dielectric report
Supply | Lead time: 2–4 weeks ex-works