## Request a quote
Request a quotation for HSC Low-Noise Triaxial
High-voltage diagnostic & low-noise cable engineer
Enter the length you need next to each size in the Sizes sheet (column A) and send this file to:
Email: info@tebaoflex.com
WhatsApp / phone: +86 185 9591 3361
Online quote form: https://tebaoflex.com/cables/low-noise-triaxial-high-voltage-cable/#template-selection
An application engineer replies within 2 business hours.

## Sizes
HSC Low-Noise Triaxial Low-Noise Triaxial High-Voltage Cable – Voltage ratings, PD level, capacitance & OD (12)
TEBAOFLEX ¦ https://tebaoflex.com/cables/low-noise-triaxial-high-voltage-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 60270 / IEEE 400.3 / MIL-C-17 Triaxial. Values for your order are confirmed in the quotation.
Enter the length you need next to each size and send this file to info@tebaoflex.com or WhatsApp +86 185 9591 3361.
Your length (m) | Order designation | HV OEM & Z0 | Center Core mm² (AWG) | Rated U_DC kV DC | PD-Free U_AC kV rms | Max PD pC | Capacitance pF/m | Overall OD mm | Weight kg/km | Series
| HSC-PE-10KV-50R | PE Low-Noise (50 Ω) | 0.95 (18 AWG) | 10 | 5 | < 0.5 | 101 | 7.4 | 82 | HSC-PE Series — Low-Loss PE/XLPE OEM Triaxial HV Cable (50 Ω / 75 Ω)
| HSC-PE-15KV-75R | PE Low-Noise (75 Ω) | 0.50 (20 AWG) | 15 | 7 | < 0.5 | 68 | 8.2 | 96 | HSC-PE Series — Low-Loss PE/XLPE OEM Triaxial HV Cable (50 Ω / 75 Ω)
| HSC-PE-20KV-50R | XLPE Low-Noise (50 Ω) | 1.50 (16 AWG) | 20 | 10 | < 0.5 | 102 | 9.6 | 134 | HSC-PE Series — Low-Loss PE/XLPE OEM Triaxial HV Cable (50 Ω / 75 Ω)
| HSC-PE-30KV-50R | XLPE Low-Noise (50 Ω) | 1.50 (16 AWG) | 30 | 14 | < 0.5 | 94 | 11.4 | 182 | HSC-PE Series — Low-Loss PE/XLPE OEM Triaxial HV Cable (50 Ω / 75 Ω)
| HSC-PE-40KV-75R | XLPE Low-Noise (75 Ω) | 1.00 (18 AWG) | 40 | 18 | < 1.0 | 72 | 13 | 228 | HSC-PE Series — Low-Loss PE/XLPE OEM Triaxial HV Cable (50 Ω / 75 Ω)
| HSC-PE-60KV-75R | XLPE Low-Noise (75 Ω) | 1.50 (16 AWG) | 60 | 25 | < 1.0 | 69 | 16.5 | 355 | HSC-PE Series — Low-Loss PE/XLPE OEM Triaxial HV Cable (50 Ω / 75 Ω)
| HSC-SIR-10KV-1.0 | Silicone Low-Noise | 1.00 (18 AWG) | 10 | 5 | < 0.5 | 148 | 7.8 | 92 | HSC-SIR Series — Flexible Silicone OEM Triaxial HV Test Cable
| HSC-SIR-20KV-1.5 | Silicone Low-Noise | 1.50 (16 AWG) | 20 | 9 | < 0.5 | 132 | 9.8 | 142 | HSC-SIR Series — Flexible Silicone OEM Triaxial HV Test Cable
| HSC-SIR-30KV-1.5 | Silicone Low-Noise | 1.50 (16 AWG) | 30 | 13 | < 1.0 | 118 | 11.8 | 198 | HSC-SIR Series — Flexible Silicone OEM Triaxial HV Test Cable
| HSC-SIR-40KV-2.5 | Silicone Low-Noise | 2.50 (14 AWG) | 40 | 17 | < 1.0 | 115 | 13.6 | 258 | HSC-SIR Series — Flexible Silicone OEM Triaxial HV Test Cable
| HSC-SIR-50KV-2.5 | Silicone Low-Noise | 2.50 (14 AWG) | 50 | 21 | < 1.0 | 108 | 15.4 | 324 | HSC-SIR Series — Flexible Silicone OEM Triaxial HV Test Cable
| HSC-SIR-60KV-2.5 | Silicone Low-Noise | 2.50 (14 AWG) | 60 | 25 | < 1.0 | 102 | 17.5 | 410 | HSC-SIR Series — Flexible Silicone OEM Triaxial HV Test Cable
Every length of HSC Low-Noise Triaxial is tested in a Faraday cage per IEC 60270 to certify partial-discharge inception/extinction and < 0.1 pC mechanical flex noise.

## Technical data
HSC Low-Noise Triaxial Low-Noise Triaxial High-Voltage Cable – Technical data
TEBAOFLEX ¦ https://tebaoflex.com/cables/low-noise-triaxial-high-voltage-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 60270 / IEEE 400.3 / MIL-C-17 Triaxial. Values for your order are confirmed in the quotation.
Group | Parameter | Value
Key data | DC voltage range | 10 to 60 kV DC (5 kV to 25 kV AC rms PD-free)
Key data | PD noise floor | < 0.5 – 1.0 pC (Exceeds IEC 60270 laboratory grade)
Key data | Triboelectric Q | ≤ 0.1 pC flex (Graphite/semicon interface damping)
Key data | Architecture | True Triaxial (HV Core + Guard Screen + EMC Shield)
Key data | Impedance | 50 Ω / 75 Ω typ. (Low-loss PE/XLPE or SiR OEM)
Key data | Inter-screen U | 2.0 kV DC (Isolated inner guard vs outer ground)
Electrical | Rated DC working voltage | 10, 15, 20, 30, 40, 50 & 60 kV DC
Electrical | Rated AC PD-free voltage | 5 kV to 25 kV AC rms (50/60 Hz)
Electrical | Partial discharge background | < 0.5 pC (up to 20 kV AC) · < 1.0 pC (25 kV AC) per IEC 60270
Electrical | Triboelectric noise (mechanically flexed) | ≤ 0.1 pC / < 50 μV peak-to-peak
Electrical | Inter-shield breakdown voltage | ≥ 2.0 kV DC (inner guard screen to outer EMC screen)
Electrical | Characteristic impedance | 50 ± 3 Ω or 75 ± 4 Ω (HSC-PE pulse series)
Thermal | Operating temperature (HSC-PE) | −30 °C to +85 °C
Thermal | Operating temperature (HSC-SIR) | −50 °C to +150 °C
Mechanical | Center conductor cross-sections | 0.50 mm², 0.75 mm², 1.0 mm², 1.5 mm², 2.5 mm²
Mechanical | Min. bend radius | 8 × OD (HSC-PE) · 6 × OD (HSC-SIR)
Resistance | Shielding attenuation (triaxial) | ≥ 95 dB (100 kHz to 1 GHz)
Resistance | Flame & oil resistance | IEC 60332-1-2 · EN 60811-404 (PUR/PVC jacket)
Standards | PD & HV test standards | IEC 60270 · IEEE 400.3 · IEC 60060-1
Standards | Connector compatibility | SHV-Triax, Redel-HV Redel Triax,Circular-HV HV & custom PD bridge plugs
Identification | Outer sheath colour | Black, Grey (RAL 7001) or Red
Identification | Surface legend | TEBAOFLEX HSC LOW-NOISE TRIAXIAL HV CABLE <1pC [kV DC]
Identification | Customs tariff (HS) | 8544.20 / 8544.60

## Construction
HSC Low-Noise Triaxial Low-Noise Triaxial High-Voltage Cable – Construction
TEBAOFLEX ¦ https://tebaoflex.com/cables/low-noise-triaxial-high-voltage-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 60270 / IEEE 400.3 / MIL-C-17 Triaxial. Values for your order are confirmed in the quotation.
No. | Layer | Material | Reference
1 | Low-noise stranded tinned / silver-plated copper core | Compacted stranded tinned or silver-plated copper center conductor (0.50 mm² to 2.5 mm², or 50 Ω / 75 Ω matched geometry)
2 | Inner & outer anti-triboelectric semiconductive layers | Co-extruded inner semiconductive stress layer and outer conductive graphite/polymer damping layer sandwiching the HV OEM
3 | Void-free HV PE/XLPE or Silicone primary OEM | Super-clean low-loss PE/XLPE (tan δ ≤ 0.0003) or flexible HV silicone rubber rated 10 kV to 60 kV DC
4 | Inner triaxial guard / signal return braid | ≥ 90 % coverage tinned copper braided inner screen in intimate contact with the outer semiconductive layer
5 | Inter-screen isolation OEM | Extruded 2.0 kV DC rated PE, TPE or silicone barrier separating the inner guard braid from the outer EMC shield
6 | Outer ≥ 90 % EMC chassis braid & protective sheath | Outer tinned copper braid (≥ 90 % coverage) under a low-friction red, grey or black PVC/PUR/Silicone outer jacket
| HSC-PE (Low-Loss PE/XLPE OEM, 50/75 Ω Pulse) | Ultra-low capacitance (67–100 pF/m) and tan δ ≤ 0.0003 for fast partial-discharge pulse propagation, PMT and proportional-counter detectors
| HSC-SIR (Extra-Flexible Silicone OEM, −50 to +150 °C) | Limp silicone HV OEM and silicone outer sheath for flexible benchtop hipot, tan-delta and electrometer test leads

## Quality and supply
HSC Low-Noise Triaxial Low-Noise Triaxial High-Voltage Cable – Testing, approvals and supply
TEBAOFLEX ¦ https://tebaoflex.com/cables/low-noise-triaxial-high-voltage-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 60270 / IEEE 400.3 / MIL-C-17 Triaxial. Values for your order are confirmed in the quotation.
Item | Detail | Frequency / scope
Routine test | Faraday-cage IEC 60270 partial-discharge test — Verify background PD < 0.5 pC (≤ 20 kV) / < 1.0 pC (25 kV AC rms) | 100 % of manufactured reels
Routine test | Core-to-guard DC hipot & guard-to-shield 2 kV DC test — 1.5 × U_DC core-to-guard for 5 min + 2.0 kV DC guard-to-outer-shield for 1 min | 100 % of manufactured reels
Sample test | Triboelectric flexing charge noise test — Measure induced charge ≤ 0.1 pC under standardised mechanical flexing | Production lot
Standard / approval | IEC 60270 < 1 pC Partial-Discharge Certified
Standard / approval | IEEE 400.3 Shielded Diagnostic Lead Compliant
Company certificate | VDE
Company certificate | UL
Company certificate | cUL
Company certificate | CE
Company certificate | KTL
Company certificate | MSHA
Company certificate | MA
Company certificate | ISO 9001
Company certificate | ISO 14001
Company certificate | ISO 45001
Sheath marking | TEBAOFLEX HSC LOW-NOISE TRIAXIAL HV CABLE <1pC IEC 60270 [kV DC]
Supply | Delivery form: Spools (50 m / 100 m) or factory-terminated triaxial test leads
Supply | Certificate: Serialized Faraday-cage PD plot & triboelectric noise certificate
Supply | Lead time: 3–4 weeks ex-works