## Request a quote
Request a quotation for HXC X-Ray
X-ray & pulsed-power HV cable engineer
Enter the length you need next to each size in the Sizes sheet (column A) and send this file to:
Email: info@tebaoflex.com
WhatsApp / phone: +86 185 9591 3361
Online quote form: https://tebaoflex.com/cables/multicore-x-ray-high-voltage-cable/#template-selection
An application engineer replies within 2 business hours.

## Sizes
HXC X-Ray Multicore X-Ray High-Voltage Cable – kV rating, filament cores, cone type & OD (12)
TEBAOFLEX ¦ https://tebaoflex.com/cables/multicore-x-ray-high-voltage-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 60522 / NEMA XR7 / IEC 60601-2-28. Values for your order are confirmed in the quotation.
Enter the length you need next to each size and send this file to info@tebaoflex.com or WhatsApp +86 185 9591 3361.
Your length (m) | Order designation | Central Filament Cores | Rated U_DC kV DC | Standard Cone Fit | Dielectric OD mm | Overall OD mm | Capacitance pF/m | Weight kg/km | Series
| HXC-XR-75KV-3C-SLIM | 3C (2×1.5 + 1×0.75 mm²) | 75 | NEMA XR7 3-Pin / Mini | 13.5 | 17.5 | 135 | 420 | Medical Diagnostic & C-Arm X-Ray HV Cables (75 kV to 100 kV DC, ±75–100 kV Bipolar)
| HXC-XR-75KV-3C-MED | 3C (2×2.0 + 1×1.0 mm²) | 75 | NEMA XR7 CA11 3-Pin | 15.2 | 19.5 | 128 | 510 | Medical Diagnostic & C-Arm X-Ray HV Cables (75 kV to 100 kV DC, ±75–100 kV Bipolar)
| HXC-XR-75KV-4C-MED | 4C (2×1.5 + 2×0.75 mm²) | 75 | NEMA XR7 CA12 4-Pin | 15.5 | 19.8 | 132 | 530 | Medical Diagnostic & C-Arm X-Ray HV Cables (75 kV to 100 kV DC, ±75–100 kV Bipolar)
| HXC-XR-90KV-3C-MED | 3C (2×2.0 + 1×1.0 mm²) | 90 | NEMA XR7 CA11 3-Pin | 16.8 | 21.2 | 120 | 595 | Medical Diagnostic & C-Arm X-Ray HV Cables (75 kV to 100 kV DC, ±75–100 kV Bipolar)
| HXC-XR-90KV-4C-MED | 4C (2×1.5 + 2×0.75 mm²) | 90 | NEMA XR7 CA12 4-Pin | 17 | 21.5 | 124 | 615 | Medical Diagnostic & C-Arm X-Ray HV Cables (75 kV to 100 kV DC, ±75–100 kV Bipolar)
| HXC-XR-100KV-4C-SIR | 4C (2×1.5 + 2×0.75 mm²) | 100 | CA12 / CT Gantry Flex | 18.2 | 23 | 138 | 690 | Medical Diagnostic & C-Arm X-Ray HV Cables (75 kV to 100 kV DC, ±75–100 kV Bipolar)
| HXC-XR-100KV-3C-R10 | 3C (2×2.5 + 1×1.5 mm²) | 100 | Euro R10 (EN 60522) | 18 | 22.8 | 118 | 680 | Industrial NDT, Security & Analytical XRD HV Cables (100 kV to 160 kV DC)
| HXC-XR-100KV-4C-R10 | 4C (2×2.5 + 2×1.0 mm²) | 100 | Euro R10 4-Pin | 18.4 | 23.2 | 122 | 705 | Industrial NDT, Security & Analytical XRD HV Cables (100 kV to 160 kV DC)
| HXC-XR-125KV-3C-R24 | 3C (2×2.5 + 1×1.5 mm²) | 125 | Euro R24 (EN 60522) | 21 | 26.2 | 115 | 890 | Industrial NDT, Security & Analytical XRD HV Cables (100 kV to 160 kV DC)
| HXC-XR-125KV-4C-R24 | 4C (2×2.5 + 2×1.0 mm²) | 125 | Euro R24 4-Pin | 21.4 | 26.6 | 118 | 920 | Industrial NDT, Security & Analytical XRD HV Cables (100 kV to 160 kV DC)
| HXC-XR-160KV-3C-R24 | 3C (2×2.5 + 1×1.5 mm²) | 160 | Euro R24 (160 kV) | 24.5 | 30 | 110 | 1180 | Industrial NDT, Security & Analytical XRD HV Cables (100 kV to 160 kV DC)
| HXC-XR-160KV-4C-R24 | 4C (2×2.5 + 2×1.0 mm²) | 160 | Euro R24 4-Pin (160 kV) | 24.8 | 30.5 | 112 | 1215 | Industrial NDT, Security & Analytical XRD HV Cables (100 kV to 160 kV DC)
Available as bulk cable on reels or complete factory-vulcanised double-ended assemblies with NEMA XR7 CA11/CA12 or Euro R10/R24 cones.

## Technical data
HXC X-Ray Multicore X-Ray High-Voltage Cable – Technical data
TEBAOFLEX ¦ https://tebaoflex.com/cables/multicore-x-ray-high-voltage-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 60522 / NEMA XR7 / IEC 60601-2-28. Values for your order are confirmed in the quotation.
Group | Parameter | Value
Key data | DC voltage range | 75 to 160 kV DC (Bipolar ±75 kV (150 kV) or unipolar)
Key data | Central cores | 3 or 4 Cores (Large/small focal-spot filament + grid)
Key data | Core-to-core U | 2.5 – 5.0 kV DC (Filament & Wehnelt grid isolation)
Key data | Capacitance | 110–145 pF/m (Low stored energy for fast kVp rise)
Key data | Cone standards | CA11 / R10 / R24 (NEMA XR7 Federal & Euro EN 60522)
Key data | Bend radius | 5–6 × OD (C-arm, CT gantry & NDT crawler flex)
Electrical | Rated DC working voltage | 75 kV, 90 kV, 100 kV, 125 kV & 160 kV DC (to ground)
Electrical | Bipolar tube system rating | 150 kVp (±75 kV) · 200 kVp (±100 kV) · 320 kVp (±160 kV)
Electrical | Filament core-to-core rating | 2.5 kV to 5.0 kV DC (up to 15 A AC/DC filament current)
Electrical | Factory DC proof test | 1.35 × to 1.5 × rated U_DC for 15 min
Electrical | Nominal capacitance | 110 pF/m to 145 pF/m
Thermal | Operating temperature (EPR/PVC/PUR) | −25 °C to +85 °C
Thermal | Operating temperature (Silicone) | −40 °C to +150 °C
Mechanical | Central conductor layout | 3-core (2×1.5/2.5 + 1×1.0) or 4-core (2×1.5/2.5 + 2×0.75)
Mechanical | Min. bend radius | 5 × OD static · 8 × OD dynamic C-arm/crane
Resistance | Disinfectant & oil resistance | Resistant to hospital cleaners & insulating oil
Resistance | Flame retardancy | IEC 60332-1-2 / UL VW-1
Standards | X-ray cable & receptacle stds | NEMA XR7 (CA11/CA12) · IEC 60522 · EN 60601-2-28
Standards | Industrial NDT cone interface | R10 (100 kV) · R24 (160/225 kV) · R30 (300 kV)
Identification | Outer sheath colour | Medical Light Grey (RAL 7035) or NDT Matte Black
Identification | Surface legend | TEBAOFLEX HXC X-RAY [kV DC] [3C/4C] NEMA XR7 / IEC 60522
Identification | Customs tariff (HS) | 8544.60 / 9022.90

## Construction
HXC X-Ray Multicore X-Ray High-Voltage Cable – Construction
TEBAOFLEX ¦ https://tebaoflex.com/cables/multicore-x-ray-high-voltage-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 60522 / NEMA XR7 / IEC 60601-2-28. Values for your order are confirmed in the quotation.
No. | Layer | Material | Reference
1 | 3 or 4 central filament & grid-bias conductors | 2 × 1.5/2.5 mm² tinned copper filament heater cores (Common, Large Focus, Small Focus) + 1 or 2 × 0.75/1.0 mm² grid/control cores insulated to 2.5–5.0 kV DC
2 | Inner semiconductive equipotential layer | Extruded semiconductive elastomer encapsulating the 3-core or 4-core central filament bundle
3 | Void-free HV EPR/HEPR or Silicone dielectric | High-dielectric-strength ethylene propylene rubber (EPR) or medical-grade silicone rubber rated 75, 100, 125 or 160 kV DC
4 | Outer semiconductive stress-control layer | Strippable semiconductive layer or conductive tape under the metallic shield
5 | ≥ 85 % tinned copper braided earth shield | Heavy tinned copper braid connected to the grounded X-ray generator tank and tube housing
6 | Medical Grey PVC / NDT Black PUR or Silicone jacket | Non-marking Medical Grey (RAL 7035/7047) PVC/Silicone for hospitals or abrasion-resistant Black PUR for industrial NDT crawlers
| HXC-XR-MED (75 / 90 / 100 kV DC Medical C-Arm & Rad) | 3-core or 4-core extra-flexible medical grey cable for bipolar ±75 kV (150 kVp) diagnostic X-ray tubes (CA11 /CA12 Federal 3/4-pin cones)
| HXC-XR-NDT (100 / 125 / 160 kV DC Industrial NDT) | 3-core or 4-core heavy-duty black PUR-jacketed cable for unipolar 160 kV or bipolar 320 kV NDT tubes (Euro R10, R24 & R30 cones)

## Quality and supply
HXC X-Ray Multicore X-Ray High-Voltage Cable – Testing, approvals and supply
TEBAOFLEX ¦ https://tebaoflex.com/cables/multicore-x-ray-high-voltage-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 60522 / NEMA XR7 / IEC 60601-2-28. Values for your order are confirmed in the quotation.
Item | Detail | Frequency / scope
Routine test | Core-to-screen HV DC proof test (105 kV to 225 kV DC) — 15-minute DC hold test between shorted central cores and grounded outer braid | 100 % of cable & assemblies
Routine test | Filament core-to-core 5.0 kV DC dielectric test — Verify inter-conductor insulation between Common, Large Focus, Small Focus & Grid cores | 100 % of cable & assemblies
Sample test | Partial discharge & capacitance uniformity test — PD ≤ 5 pC and capacitance within ± 5 % of nominal | Production lot
Standard / approval | NEMA XR7 & IEC 60522 X-Ray HV Cable Compliant
Standard / approval | IEC 60601-2-28 Medical X-Ray Assembly Compatible
Company certificate | VDE
Company certificate | UL
Company certificate | cUL
Company certificate | CE
Company certificate | KTL
Company certificate | MSHA
Company certificate | MA
Company certificate | ISO 9001
Company certificate | ISO 14001
Company certificate | ISO 45001
Sheath marking | TEBAOFLEX HXC X-RAY HIGH VOLTAGE CABLE [kV DC] NEMA XR7 / IEC 60522
Supply | Delivery form: Bulk reels (50 m / 100 m) or terminated HV cone assemblies (2 m to 25 m)
Supply | Terminations: Federal Standard CA11 (3-pin), CA12 (4-pin), Euro R10, R24 & pancake plugs
Supply | Lead time: 3–4 weeks ex-works