## Request a quote
Request a quotation for LEAD108B / LEAD108C
Medical & scientific instrumentation cable engineer
Enter the length you need next to each size in the Sizes sheet (column A) and send this file to:
Email: info@tebaoflex.com
WhatsApp / phone: +86 185 9591 3361
Online quote form: https://tebaoflex.com/cables/mri-conditional-carbon-composition-biopotential-lead-assembly/#footer-enquiry-section
An application engineer replies within 2 business hours.

## Sizes
LEAD108B / LEAD108C MRI-Conditional Carbon Biopotential Lead Assembly – Lead styles, lengths & DC resistance (12)
TEBAOFLEX ¦ https://tebaoflex.com/cables/mri-conditional-carbon-composition-biopotential-lead-assembly/ ¦ Data revision 2026-10-21
Typical values for this cable type, based on ASTM F2503 /ASTM F2182 / IEC 60601-2-33 / ANSI/AAMI EC53. Values for your order are confirmed in the quotation.
Enter the length you need next to each size and send this file to info@tebaoflex.com or WhatsApp +86 185 9591 3361.
Your length (m) | Order designation | Lead Architecture & Shielding | Electrode End | Length m | Unit R Ω/m | Total R Ω | Cable OD mm | MR Field Rating | Series
| MRC-LEAD108B-SNAP-060 | Unshielded 3K Carbon Core | Carbon Snap (3.9 mm) | 0.6 | 110 | 66 | 2.3 | 1.5T / 3T / 7T | LEAD108B Series — Unshielded Carbon-Composition Single Leads
| MRC-LEAD108B-SNAP-100 | Unshielded 3K Carbon Core | Carbon Snap (3.9 mm) | 1 | 110 | 110 | 2.3 | 1.5T / 3T / 7T | LEAD108B Series — Unshielded Carbon-Composition Single Leads
| MRC-LEAD108B-SNAP-150 | Unshielded 3K Carbon Core (LEAD108B std) | Carbon Snap (3.9 mm) | 1.5 | 110 | 165 | 2.3 | 1.5T / 3T / 7T | LEAD108B Series — Unshielded Carbon-Composition Single Leads
| MRC-LEAD108B-SNAP-200 | Unshielded 3K Carbon Core | Carbon Snap (3.9 mm) | 2 | 110 | 220 | 2.3 | 1.5T / 3T | LEAD108B Series — Unshielded Carbon-Composition Single Leads
| MRC-LEAD108B-CLIP-150 | Unshielded 3K Carbon Core | Radiolucent Pinch Clip | 1.5 | 110 | 165 | 2.3 | 1.5T / 3T / 7T | LEAD108B Series — Unshielded Carbon-Composition Single Leads
| MRC-LEAD108B-HR-150 | High-Damping 300 Ω/m Carbon Core | Carbon Snap (3.9 mm) | 1.5 | 300 | 450 | 2.3 | 3T / 7T / 9.4T | LEAD108B Series — Unshielded Carbon-Composition Single Leads
| MRC-LEAD108C-SCR-100 | Carbon-Polymer Screened Coax | Carbon Snap (3.9 mm) | 1 | 120 | 120 | 3.2 | 1.5T / 3T / 7T | LEAD108C & Twisted Series — Carbon-Screened & Differential Pair Leads
| MRC-LEAD108C-SCR-150 | Carbon-Polymer Screened Coax (LEAD108C std) | Carbon Snap (3.9 mm) | 1.5 | 120 | 180 | 3.2 | 1.5T / 3T / 7T | LEAD108C & Twisted Series — Carbon-Screened & Differential Pair Leads
| MRC-LEAD108C-SCR-200 | Carbon-Polymer Screened Coax | Carbon Snap (3.9 mm) | 2 | 120 | 240 | 3.2 | 1.5T / 3T | LEAD108C & Twisted Series — Carbon-Screened & Differential Pair Leads
| MRC-LEAD108-TP2-150 | 2-Lead Twisted Carbon Pair (Bipolar) | 2 × Carbon Snaps | 1.5 | 110 | 165 | 4.4 | 1.5T / 3T / 7T | LEAD108C & Twisted Series — Carbon-Screened & Differential Pair Leads
| MRC-LEAD108-TP3-150 | 3-Lead Twisted Carbon Triad (+ / − / GND) | 3 × Carbon Snaps (R/W/B) | 1.5 | 110 | 165 | 5 | 1.5T / 3T / 7T | LEAD108C & Twisted Series — Carbon-Screened & Differential Pair Leads
| MRC-LEAD108-TP3-300 | 3-Lead Twisted Carbon Triad (+ / − / GND) | 3 × Carbon Snaps (R/W/B) | 3 | 110 | 330 | 5 | 1.5T / 3T | LEAD108C & Twisted Series — Carbon-Screened & Differential Pair Leads
To minimise gradient-switching loop pickup and RF heating inside the MRI bore, always route leads straight down the centre of the bore without coiling and use twisted-pair configurations where possible.

## Technical data
LEAD108B / LEAD108C MRI-Conditional Carbon Biopotential Lead Assembly – Technical data
TEBAOFLEX ¦ https://tebaoflex.com/cables/mri-conditional-carbon-composition-biopotential-lead-assembly/ ¦ Data revision 2026-10-21
Typical values for this cable type, based on ASTM F2503 /ASTM F2182 / IEC 60601-2-33 / ANSI/AAMI EC53. Values for your order are confirmed in the quotation.
Group | Parameter | Value
Key data | MR field rating | 1.5T / 3T / 7T (MR-Conditional per ASTM F2503)
Key data | Core material | 100 % Carbon (Zero ferromagnetic metal inside bore)
Key data | Distributed R | 90–350 Ω/m (Damps RF standing-wave heating)
Key data | Electrode end | Carbon Snap/Clip (Radiotransparent Ag/AgCl snap or pinch)
Key data | Amplifier end | 1.5 mm DIN (Touch-proof IEC 60601-1 / DIN 42802)
Key data | Jacket | USP Class VI (Biocompatible medical TPU / silicone)
Electrical | Distributed DC resistance | 100 Ω/m ± 20 % (standard) or 300 Ω/m ± 20 % (high-damping)
Electrical | Total lead resistance | 100 Ω to 600 Ω (negligible vs ≥ 10 MΩ amplifier input)
Electrical | Dielectric strength | 1,500 V AC rms for 1 min (IEC 60601-1)
Electrical | RF heating (ASTM F2182) | ΔT < 1.5 °C at 4 W/kg whole-body SAR (1.5 T & 3.0 T)
Electrical | Image artefact (ASTM F2119) | Zero metallic blooming / < 1 mm spin-echo & GRE void
Thermal | Clinical operating range | +10 °C to +45 °C
Thermal | Storage & cleaning temperature | −20 °C to +60 °C
Mechanical | Cable outer diameter | 2.3 mm (single unshielded) · 3.4 mm (carbon-screened / twisted)
Mechanical | Electrode termination | All-carbon/polymer 3.9 mm snap or radiolucent micro-clip
Mechanical | Amplifier termination | 1.5 mm female touch-proof safety socket (DIN 42802)
Mechanical | Flex life (ANSI/AAMI EC53) | ≥ 25,000 swing cycles at strain relief
Mechanical | Tensile pull strength | ≥ 35 N termination pull-off force
Resistance | Biocompatibility | ISO 10993-5 (cytotoxicity) & ISO 10993-10 (sensitization)
Resistance | Disinfectant resistance | 70 % IPA, Quaternary Ammonium, hospital surface wipes & 3 % H2O2
Resistance | Latex & phthalate status | 100 % natural-rubber-latex-free & DEHP-free
Standards | MR safety standards | ASTM F2503 · ASTM F2182 · ASTM F2119 · IEC 60601-2-33
Standards | ECG trunk/lead performance | ANSI/AAMI EC53 / IEC 60601-2-27
Standards | Safety connector standard | DIN 42802 / IEC 60601-1 touch-proof 1.5 mm
Identification | Jacket colour | Matte Charcoal Black with colour-coded AHA/IEC rings
Identification | MR marking | Permanent yellow 'MR CONDITIONAL' tag with length & R code
Identification | Customs tariff (HS) | 9018.11 / 8544.42

## Construction
LEAD108B / LEAD108C MRI-Conditional Carbon Biopotential Lead Assembly – Construction
TEBAOFLEX ¦ https://tebaoflex.com/cables/mri-conditional-carbon-composition-biopotential-lead-assembly/ ¦ Data revision 2026-10-21
Typical values for this cable type, based on ASTM F2503 /ASTM F2182 / IEC 60601-2-33 / ANSI/AAMI EC53. Values for your order are confirmed in the quotation.
No. | Layer | Material | Reference
1 | Multi-filament carbon-composition core | 3,000 to 6,000 continuous high-purity polyacrylonitrile (PAN) carbon filaments (distributed resistance 90–350 Ω/m, magnetic susceptibility χ ≈ −10⁻⁵)
2 | Medical-grade TPE / fluoroelastomer primary insulation | Flexible dielectric core insulation (2.0 kV dielectric withstand)
| LEAD108B-EQ (Unshielded Carbon Lead, 1.0–1.5 m) | Single-core 2.3 mm OD carbon-composition lead with radiolucent carbon snap and 1.5 mm touch-proof DIN 42802 socket
| LEAD108C-EQ (Carbon-Screened Lead / Twisted Pair) | Carbon-polymer screened or tightly twisted dual-carbon differential pair for minimal gradient artefact in 3T/7T EPI sequences

## Quality and supply
LEAD108B / LEAD108C MRI-Conditional Carbon Biopotential Lead Assembly – Testing, approvals and supply
TEBAOFLEX ¦ https://tebaoflex.com/cables/mri-conditional-carbon-composition-biopotential-lead-assembly/ ¦ Data revision 2026-10-21
Typical values for this cable type, based on ASTM F2503 /ASTM F2182 / IEC 60601-2-33 / ANSI/AAMI EC53. Values for your order are confirmed in the quotation.
Item | Detail | Frequency / scope
Routine test | 100 % DC end-to-end carbon resistance & continuity check — Verify distributed resistance within ± 15 % of nominal specification | Every lead assembly
Routine test | 1,500 V AC dielectric withstand test — Zero breakdown through medical TPU jacket for 60 s | Every production lot
Type test | ASTM F2182 phantom RF heating test at 64 MHz (1.5T) & 128 MHz (3T) — Temperature rise ΔT < 1.5 °C at electrode contact in gel phantom | Type qualification
Type test | ASTM F2119 gradient-echo & spin-echo susceptibility artefact test — Zero ferromagnetic distortion on 3.0 T EPI sequences | Type qualification
Standard / approval | ASTM F2503 / ASTM F2182 MR-Conditional Verified
Standard / approval | ANSI/AAMI EC53 & ISO 10993 Biocompatible
Company certificate | VDE
Company certificate | UL
Company certificate | cUL
Company certificate | CE
Company certificate | KTL
Company certificate | MSHA
Company certificate | MA
Company certificate | ISO 9001
Company certificate | ISO 14001
Company certificate | ISO 45001
Sheath marking | TEBAOFLEX MRC-LEAD108 MR CONDITIONAL CARBON LEAD DIN 42802
Supply | Packaging: Clean medical pouch in colour-coded sets of 3, 5 or 10 leads
Supply | Colour coding: AHA (Red/White/Black/Green/Brown) or IEC (Red/Yellow/Black/Green)
Supply | Amplifier interface: Standard 1.5 mm DIN 42802 female safety socket or multi-pin trunk
Supply | Lead time: 2–3 weeks ex-works