## Request a quote
Request a quotation for Low-Noise Cable for Piezoelectric Sensors
Application engineer · vibration, acoustic & aerospace cables
Enter the length you need next to each size in the Sizes sheet (column A) and send this file to:
Email: info@tebaoflex.com
WhatsApp / phone: +86 185 9591 3361
Online quote form: https://tebaoflex.com/cables/low-noise-cable-for-piezoelectric-sensors/#template-selection
An application engineer replies within 2 business hours.

## Sizes
Low-Noise Cable for Piezoelectric Sensors – Dimensions, triboelectric noise & temperature ratings (12)
TEBAOFLEX ¦ https://tebaoflex.com/cables/low-noise-cable-for-piezoelectric-sensors/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEST-RP-DTE011 / MIL-DTL-17 / ISO 16063. Values for your order are confirmed in the quotation.
Enter the length you need next to each size and send this file to info@tebaoflex.com or WhatsApp +86 185 9591 3361.
Your length (m) | Order designation | Construction | Dielectric / Jacket | Noise & temp. rating | Outer Ø ≈ mm | Weight ≈ kg/km | Min. bend R mm | Series
| PIEZO-LN-50-195 | 50 Ω Coax · AWG 28 (7×36 SPC) | PTFE + Carbon / FEP | ≤ 15 fC · −65…+200 °C | 1.95 | 8.5 | 10 | 50 Ω Low-Noise Coaxial Charge-Mode Cables (10-32 UNF / BNC)
| PIEZO-LN-50-250 | 50 Ω Coax · AWG 26 (7×34 SPC) | PTFE + Carbon / FEP | ≤ 10 fC · −65…+200 °C | 2.5 | 13 | 12 | 50 Ω Low-Noise Coaxial Charge-Mode Cables (10-32 UNF / BNC)
| PIEZO-LN-50-280 (RG-174 LN) | 50 Ω Coax · AWG 26 (7×34 SPC) | PTFE + Carbon / PFA | ≤ 5 fC · −70…+260 °C | 2.8 | 16.5 | 14 | 50 Ω Low-Noise Coaxial Charge-Mode Cables (10-32 UNF / BNC)
| PIEZO-LN-50-320 | 50 Ω Coax · AWG 24 (19×36 SPC) | PTFE + Carbon / FEP | ≤ 10 fC · −65…+200 °C | 3.2 | 21 | 16 | 50 Ω Low-Noise Coaxial Charge-Mode Cables (10-32 UNF / BNC)
| PIEZO-LN-PUR-300 | 50 Ω Coax · AWG 26 (19×38 Cu) | PE + Carbon / PUR | ≤ 25 fC · −40…+105 °C | 3 | 14.5 | 15 | 50 Ω Low-Noise Coaxial Charge-Mode Cables (10-32 UNF / BNC)
| PIEZO-LN-TP 1×2×AWG26 | 1 Pair × AWG 26 (19×38 SPC) | PTFE + Carbon / FEP | Diff. Charge · −65…+200 °C | 3.4 | 19.5 | 17 | Differential Twisted-Pair Low-Noise Charge-Mode Cables (Turbine / Industrial)
| PIEZO-LN-TP 1×2×AWG24 | 1 Pair × AWG 24 (19×36 SPC) | PTFE + Carbon / FEP | Diff. Charge · −65…+200 °C | 4.1 | 28 | 20 | Differential Twisted-Pair Low-Noise Charge-Mode Cables (Turbine / Industrial)
| PIEZO-LN-TP 1×2×AWG22 | 1 Pair × AWG 22 (19×34 SPC) | PTFE + Carbon / PFA | Diff. Charge · −70…+260 °C | 4.8 | 39 | 24 | Differential Twisted-Pair Low-Noise Charge-Mode Cables (Turbine / Industrial)
| PIEZO-LN-TP 2×2×AWG24 | 2 Pairs × AWG 24 (19×36 SPC) | PTFE + Carbon / FEP | Dual-Axis · −65…+200 °C | 6.4 | 62 | 32 | Differential Twisted-Pair Low-Noise Charge-Mode Cables (Turbine / Industrial)
| PIEZO-LN-TRIAX-280 | 50 Ω Triax · AWG 26 SPC | PTFE + Carbon / FEP | Driven Guard · −65…+200 °C | 2.85 | 19 | 14 | Concentric Triaxial & 4-Wire Triaxial IEPE Sensor Cables
| PIEZO-LN-TRIAX-360 | 50 Ω Triax · AWG 24 SPC | PTFE + Carbon / PFA | Driven Guard · −70…+260 °C | 3.6 | 29 | 18 | Concentric Triaxial & 4-Wire Triaxial IEPE Sensor Cables
| PIEZO-IEPE-4×AWG26 | 4-Core × AWG 26 SPC | FEP / FEP Braided | Triaxial IEPE · −65…+200 °C | 3.5 | 23 | 18 | Concentric Triaxial & 4-Wire Triaxial IEPE Sensor Cables
When terminating low-noise coaxial cable, strip back or clean any loose conductive carbon residue from the cut end of the PTFE dielectric so insulation resistance stays ≥ 10¹² Ω.

## Electrical parameters by cable
Low-Noise Cable for Piezoelectric Sensors – Electrical parameters by cable geometry (4)
TEBAOFLEX ¦ https://tebaoflex.com/cables/low-noise-cable-for-piezoelectric-sensors/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEST-RP-DTE011 / MIL-DTL-17 / ISO 16063. Values for your order are confirmed in the quotation.
Geometry | Impedance Ω | Nom. C pF/m | Max. tribo noise fC pk-pk | Min. IR @ 20 °C GΩ·m
50 Ω Coax (PTFE + Carbon / FEP or PFA) | 50 ± 2 | 96 | 10 | 10000
50 Ω Coax (PE + Carbon / PUR) | 50 ± 3 | 100 | 25 | 5000
Differential Twisted Pair (PTFE + Carbon) | 100 ± 10 | 85 | 15 | 10000
50 Ω Concentric Triaxial (PTFE + Carbon) | 50 ± 2 | 96 | 10 | 10000
1 fC (femtocoulomb) = 0.001 pC (picocoulomb). Untreated RG-174 coaxial cable can generate > 5,000 fC (5 pC) under the same vibration.

## Technical data
Low-Noise Cable for Piezoelectric Sensors – Technical data
TEBAOFLEX ¦ https://tebaoflex.com/cables/low-noise-cable-for-piezoelectric-sensors/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEST-RP-DTE011 / MIL-DTL-17 / ISO 16063. Values for your order are confirmed in the quotation.
Group | Parameter | Value
Key data | Tribo-noise | ≤ 5–15 fC (0.005–0.015 pC per IEST whip)
Key data | Insulation R | ≥ 10¹³ Ω·m (prevents charge-amp drift)
Key data | Temperature | −70…+260 °C (PFA · −65…+200 °C FEP · PUR)
Key data | Capacitance | 95–100 pF/m (50 Ω coax · ~85 pF/m diff. pair)
Key data | Screen coverage | ≥ 95 % SPC Braid (silver-plated copper optical cov.)
Key data | Bend radius | 5 × OD (sub-miniature Ø 1.95–4.80 mm)
Electrical | Max. triboelectric charge noise | ≤ 5 fC to ≤ 15 fC (0.005–0.015 pC) peak-to-peak
Electrical | Insulation resistance | ≥ 10,000 GΩ·m (10¹³ Ω·m) at +20 °C
Electrical | Characteristic impedance (coax / triax) | 50 Ω ± 2 Ω
Electrical | Nominal capacitance | 95–100 pF/m (50 Ω coax) · ~85 pF/m (differential pair)
Electrical | Operating & test voltage | 250 V rms working · 1000 V rms AC dielectric test
Thermal | Operating temperature (PFA jacket) | −70 °C to +260 °C
Thermal | Operating temperature (FEP jacket) | −65 °C to +200 °C
Thermal | Operating temperature (PUR jacket) | −40 °C to +105 °C
Mechanical | Min. bending radius | 5 × OD (dynamic vibration) · 3 × OD (static routing)
Mechanical | Connector compatibility | 10-32 UNF Microdot, M5, BNC, TNC, LEMO 00/0B, MIL-C-5015
Resistance | Fluid & chemical resistance | Resistant to Jet-A1, Skydrol, synthetic lubricants & solvents
Resistance | Flame retardance | UL 94 V-0 / IEC 60332-1-2 (FEP & PFA fluoropolymer jackets)
Standards | Triboelectric test method | IEST-RP-DTE011 & Perls / ISA S37.10 shaker whip test
Standards | Coaxial construction basis | MIL-DTL-17 (RG-174 / RG-178 / RG-316 low-noise derivatives)
Identification | Jacket colour | Translucent Blue, Salmon/Orange, Black or White (by temp grade)
Identification | Customs tariff (HS) | 8544.20 / 8544.49

## Construction
Low-Noise Cable for Piezoelectric Sensors – Construction
TEBAOFLEX ¦ https://tebaoflex.com/cables/low-noise-cable-for-piezoelectric-sensors/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEST-RP-DTE011 / MIL-DTL-17 / ISO 16063. Values for your order are confirmed in the quotation.
No. | Layer | Material | Reference
1 | Fine-stranded silver-plated copper (SPC / SPCCS) | 7-strand or 19-strand silver-plated oxygen-free copper or high-strength copper-clad steel (AWG 28 to AWG 22)
2 | Virgin extruded PTFE dielectric | High-purity polytetrafluoroethylene (PTFE) with insulation resistance ≥ 10¹³ Ω·m
3 | Semi-conductive carbon/graphite-PTFE triboelectric layer | Sintered conductive carbon-loaded PTFE dispersion layer bonded directly to the outer surface of the PTFE dielectric
4 | High-density silver-plated copper braid screen | Tightly woven silver-plated copper (SPC) braid with ≥ 95 % optical coverage in intimate contact with the conductive layer
5 | High-temperature FEP, PFA or flexible PUR jacket | Extruded fluoropolymer FEP (−65…+200 °C), PFA (−70…+260 °C), or oil-resistant PUR (−40…+105 °C)
| PIEZO-LN Coaxial (50 Ω) | Sub-miniature Ø 1.95, 2.50, 2.80 (RG-174 LN) and 3.20 mm single-ended charge cable for 10-32 UNF / M5 / BNC
| PIEZO-LN-TP Differential Pair | Individually tribo-treated twisted pair + overall SPC braid for balanced differential charge amplifiers (2-pin MIL / Lemo)
| PIEZO-LN-TRIAX (Concentric Triaxial) | Inner signal + tribo layer + inner guard braid + outer ground braid for driven-guard charge systems

## Quality and supply
Low-Noise Cable for Piezoelectric Sensors – Testing, approvals and supply
TEBAOFLEX ¦ https://tebaoflex.com/cables/low-noise-cable-for-piezoelectric-sensors/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEST-RP-DTE011 / MIL-DTL-17 / ISO 16063. Values for your order are confirmed in the quotation.
Item | Detail | Frequency / scope
Routine test | IEST-RP-DTE011 dynamic shaker triboelectric whip test — Generated charge ≤ 5–15 fC peak-to-peak under standard mechanical displacement | Every production lot
Routine test | High-resistance electrometer insulation test (≥ 10¹² Ω) — Verify zero carbon bridging and ≥ 10,000 GΩ·m dielectric resistance | Every manufactured reel
Standard / approval | IEST-RP-DTE011 Low-Triboelectric-Noise Qualified
Standard / approval | MIL-DTL-17 Fluoropolymer Coaxial Basis
Standard / approval | CE, RoHS 2 (2011/65/EU) & REACH
Company certificate | VDE
Company certificate | UL
Company certificate | cUL
Company certificate | CE
Company certificate | KTL
Company certificate | MSHA
Company certificate | MA
Company certificate | ISO 9001
Company certificate | ISO 14001
Company certificate | ISO 45001
Sheath marking | TEBAOFLEX PIEZO-LN LOW NOISE SENSOR CABLE
Supply | Delivery form: Non-returnable fumigated wooden drums
Supply | Standard lengths: 500 m · 1000 m standard drums (cut lengths on request)
Supply | MOQ: 500 m (stock sizes) · 1000 m (custom)
Supply | Lead time: 2–4 weeks ex-works
Supply | Incoterms: EXW, FCA, FOB, CIF, DAP