## Request a quote
Request a quotation for MI Triaxial Signal Cable
Application engineer · instrumentation & sensing cables
Enter the length you need next to each size in the Sizes sheet (column A) and send this file to:
Email: info@tebaoflex.com
WhatsApp / phone: +86 185 9591 3361
Online quote form: https://tebaoflex.com/cables/mineral-insulated-metal-sheathed-low-noise-triaxial-in-core-detector-cable/#template-selection
An application engineer replies within 2 business hours.

## Sizes
MI Triaxial Signal Cable – Dimensions & triaxial electrical data (4)
TEBAOFLEX ¦ https://tebaoflex.com/cables/mineral-insulated-metal-sheathed-low-noise-triaxial-in-core-detector-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on ASTM E844 / IEC 60702-1 / IEC 60568. Values for your order are confirmed in the quotation.
Enter the length you need next to each size and send this file to info@tebaoflex.com or WhatsApp +86 185 9591 3361.
Your length (m) | Order designation | Outer / Inner Tube / Core | Outer Ø mm | Weight ≈ kg/km | Core-Guard C pF/m | Core-Guard V V AC | Series
| MI-TRIAX-3.0-INC | Inconel 600 / Inconel 600 / Ni (MgO) | 3 | 46 | 155 | 1000 | Mineral-Insulated Metal-Sheathed (MIMS) Triaxial Cables
| MI-TRIAX-4.0-INC | Inconel 600 / OFHC Cu / OFHC Cu (MgO) | 4 | 82 | 135 | 1200 | Mineral-Insulated Metal-Sheathed (MIMS) Triaxial Cables
| MI-TRIAX-4.76-INC (.188") | Inconel 600 / OFHC Cu / OFHC Cu (50 Ω) | 4.76 | 115 | 125 | 1500 | Mineral-Insulated Metal-Sheathed (MIMS) Triaxial Cables
| MI-TRIAX-6.35-304L (.250") | AISI 304L / OFHC Cu / OFHC Cu (50 Ω) | 6.35 | 198 | 115 | 1800 | Mineral-Insulated Metal-Sheathed (MIMS) Triaxial Cables
Both MgO insulation layers are simultaneously vacuum-baked and tested to ≥ 10¹² Ω·m before hermetic end sealing.

## Technical data
MI Triaxial Signal Cable – Technical data
TEBAOFLEX ¦ https://tebaoflex.com/cables/mineral-insulated-metal-sheathed-low-noise-triaxial-in-core-detector-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on ASTM E844 / IEC 60702-1 / IEC 60568. Values for your order are confirmed in the quotation.
Group | Parameter | Value
Key data | Architecture | 3 Concentric (Core + inner guard + outer sheath)
Key data | Dielectric | Dual 99.4 % MgO (Core-to-guard & guard-to-sheath)
Key data | Insul. resist. | ≥ 10¹² Ω·m (≥ 10⁸ Ω·m @ +600 °C)
Key data | Operating temp. | -269 to +600 °C (Up to +850 °C Inconel peak)
Key data | Outer diameter | 3.0 to 6.35 mm (0.118" to 0.250" triaxial ODs)
Key data | Leakage cancel | Driven guard rin (Eliminates HV polarisation leakage)
Electrical | Characteristic impedance | 50 ± 4 Ω (core-to-inner-shield) or low-capacitance charge geometry
Electrical | Insulation resistivity | ≥ 10¹² Ω·m (@ 20 °C) · ≥ 10⁸ Ω·m (@ +600 °C) both layers
Electrical | Core-to-guard capacitance | 110 to 165 pF/m (@ 1 kHz)
Electrical | Dielectric test voltage | 1,000–1,500 V AC (core-guard) · 1,000 V AC (guard-sheath)
Thermal | Continuous operating temp. | -269 °C to +600 °C (up to +850 °C peak for Inconel 600)
Thermal | Thermal cycling stability | Zero dielectric cracking across 100 thermal cycles (+20 ↔ +600 °C)
Thermal | Fire & outgassing | 100 % inorganic; zero outgassing or combustion
Mechanical | Min. cold bend radius | 4 × OD (fixed forming) · 6 × OD (recommended layout)
Mechanical | Coolant pressure rating | 25 MPa (250 bar) external hydrostatic pressure at +350 °C
Mechanical | Helium leak rate | < 1 × 10⁻⁹ mbar·L/s (100 % mass-spectrometer tested)
Resistance | Neutron & gamma fluence | > 10¹⁰ Gy gamma and > 10²² n/cm² neutron fluence
Resistance | Transfer impedance (triax) | < 1 mΩ/m from 10 kHz to 30 MHz (solid tubular double shield)
Resistance | Low cobalt purity | Co ≤ 0.05 % (500 ppm) nuclear grade alloy
Standards | Nuclear MIMS standard | ASTM E844, IEC 60702-1 & IEC 60568 (In-Core Instrumentation)
Standards | Quality assurance | ISO 19443 / ASME NQA-1 / 10 CFR 50 App B
Standards | RoHS & REACH | 2011/65/EU & EC 1907/2006 compliant
Identification | Sheath finish | Bright annealed seamless Inconel 600 or AISI 304L/316L tube
Identification | Customs tariff (HS) | 8544.2000
Identification | Traceability | Laser-etched heat lot tag with CMTR chemical & boron/cobalt assay

## Construction
MI Triaxial Signal Cable – Construction
TEBAOFLEX ¦ https://tebaoflex.com/cables/mineral-insulated-metal-sheathed-low-noise-triaxial-in-core-detector-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on ASTM E844 / IEC 60702-1 / IEC 60568. Values for your order are confirmed in the quotation.
No. | Layer | Material | Reference
1 | Solid centre signal conductor | Solid OFHC copper, Nickel 200 or Inconel 600 centre rod (0.40 to 1.0 mm diameter)
2 | Inner high-purity compacted MgO dielectric | Nuclear-grade ≥ 99.4 % Magnesium Oxide (MgO, low boron/cobalt) compacted around the centre conductor
3 | Seamless concentric inner guard / HV return tube | Seamless drawn OFHC copper, Nickel or Inconel 600 intermediate tubular shield (100 % optical coverage)
4 | Outer compacted MgO isolation barrier | Second concentric layer of high-density 99.4 % MgO mineral insulation
5 | Seamless Inconel 600 or AISI 304L outer pressure sheath | Cold-drawn, bright-annealed Alloy 600 (Inconel 600) or low-cobalt AISI 304L/316L outer tube
| MI-TRIAX-3.0-INC (3.00 mm · .118 in) | Compact 3.0 mm Inconel 600 triaxial cable for in-core fission chambers and SPND thimbles
| MI-TRIAX-4.76-INC (4.76 mm · .188 in) | Standard 3/16" Inconel 600 triaxial cable with low-loss copper inner core & guard tube
| MI-TRIAX-6.35-304L (6.35 mm · .250 in) | Heavy-duty 1/4" 50 Ω triaxial cable for ex-core wide-range neutron channels

## Quality and supply
MI Triaxial Signal Cable – Testing, approvals and supply
TEBAOFLEX ¦ https://tebaoflex.com/cables/mineral-insulated-metal-sheathed-low-noise-triaxial-in-core-detector-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on ASTM E844 / IEC 60702-1 / IEC 60568. Values for your order are confirmed in the quotation.
Item | Detail | Frequency / scope
Routine test | 100 % dual-layer IR (@ 500 V DC), AC hi-pot & TDR check — Core-to-guard and guard-to-sheath IR ≥ 10¹² Ω·m @ 20 °C | Every manufactured coil
Routine test | 100 % Helium mass-spectrometer leak & 25 MPa hydrotest — Leak rate < 1 × 10⁻⁹ mbar·L/s and zero sheath deformation | Every manufactured coil
Type test | +600 °C guard-ring leakage current & breakdown test — Maintains ≥ 10⁸ Ω·m on both concentric MgO walls at +600 °C | Per production lot
Standard / approval | ASTM E844 & IEC 60568
Standard / approval | IEC 60702-1 MIMS Standard
Standard / approval | ISO 19443 / ASME NQA-1 QA
Company certificate | VDE
Company certificate | UL
Company certificate | cUL
Company certificate | CE
Company certificate | KTL
Company certificate | MSHA
Company certificate | MA
Company certificate | ISO 9001
Company certificate | ISO 14001
Company certificate | ISO 45001
Sheath marking | Laser-etched tag: TEBAOFLEX MI-TRIAX INCONEL 600 / MgO ASTM E844
Supply | Delivery form: Fumigated wooden drums, steel reels or precision spools
Supply | Standard lengths: 500 m · 1000 m (or custom cut/spooled lengths)
Supply | MOQ: 100 m (standard) · 500 m (custom)
Supply | Lead time: 2–4 weeks ex-works
Supply | Incoterms: EXW, FOB, CIF, DAP