## Request a quote
Request a quotation for Micro-Coaxial Cable for Medical and Precision Equipment
Specify AWG gauge (36–46 AWG), impedance (50 Ω or 75 Ω), maximum pF/m capacitance, single-coax bobbins or channel count (16–256 ch), and outer jacket material.
Enter the length you need next to each size in the Sizes sheet (column A) and send this file to:
Email: info@tebaoflex.com
WhatsApp / phone: +86 185 9591 3361
Online quote form: https://tebaoflex.com/cables/micro-coaxial-cable-for-medical-and-precision-equipment/#template-selection
An application engineer replies within 2 business hours.

## Sizes
Micro-Coaxial Cable for Medical and Precision Equipment – Dimensions & construction (4)
TEBAOFLEX ¦ https://tebaoflex.com/cables/micro-coaxial-cable-for-medical-and-precision-equipment/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 61196-1 / ISO 10993 / UL Style 1354. Values for your order are confirmed in the quotation.
Enter the length you need next to each size and send this file to info@tebaoflex.com or WhatsApp +86 185 9591 3361.
Your length (m) | Order designation | Construction / Variant | Conductor | Outer Ø ≈ mm | Metal kg/km | Weight ≈ kg/km | Min. bend R fixed mm | Series
| 38 AWG 50 Ω Solid PFA Micro-Coax | 50 Ω Spiral Ag-Cu Shield (PFA) | 38 AWG (7×0.040) | 0.44 | 0.4 | 0.7 | 2 | Standard Micro-Coaxial Cable for Medical and Precision Equipment variants
| 40 AWG 75 Ω Low-Cap Foamed PFA (48 pF/m) | 75 Ω Spiral Ag-Cu Shield (PFA) | 40 AWG (7×0.032) | 0.36 | 0.3 | 0.5 | 1 | Standard Micro-Coaxial Cable for Medical and Precision Equipment variants
| 42 AWG 75 Ω Low-Cap Foamed PFA (45 pF/m) | 75 Ω Spiral Ag-Cu Shield (PFA) | 42 AWG (7×0.025) | 0.3 | 0.2 | 0.3 | 1 | Standard Micro-Coaxial Cable for Medical and Precision Equipment variants
| 128-Channel 40 AWG Ultrasound Probe Bundle | 128× 40 AWG Coax + Overall Braid | 128 × 40 AWG | 8.2 | 46 | 88 | 33 | Standard Micro-Coaxial Cable for Medical and Precision Equipment variants
Dimensions and weights are nominal production values subject to manufacturing tolerances per IEC 61196-1 / ISO 10993 / UL Style 1354.

## Attenuation & RF power handling
Micro-Coaxial Cable for Medical and Precision Equipment – Attenuation & RF power handling (6)
TEBAOFLEX ¦ https://tebaoflex.com/cables/micro-coaxial-cable-for-medical-and-precision-equipment/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 61196-1 / ISO 10993 / UL Style 1354. Values for your order are confirmed in the quotation.
Frequency | Nominal attenuation dB / 100 m | Nominal attenuation dB / 100 ft | Max. CW RF power W (@ 40 °C) | Impedance Z₀ Ω | Velocity VoP %
5 MHz (Abdominal Ultrasound) | 18 | 5.49 | 15 | 50 | 78
15 MHz (Vascular / Linear Probe) | 31 | 9.45 | 10 | 50 | 78
40 MHz (High-Freq IVUS / Derm) | 52 | 15.85 | 6 | 50 | 78
100 MHz (Endoscope LVDS / Clock) | 82 | 24.99 | 4 | 50 | 78
500 MHz (High-Speed Camera / ATE) | 185 | 56.39 | 2 | 50 | 78
1000 MHz (1 GHz Micro-Interconnect) | 265 | 80.77 | 1 | 50 | 78
Attenuation values apply at +20 °C ambient into a matched 50 Ω load; CW power ratings assume VSWR 1.0:1 at +40 °C ambient.

## Technical data
Micro-Coaxial Cable for Medical and Precision Equipment – Technical data
TEBAOFLEX ¦ https://tebaoflex.com/cables/micro-coaxial-cable-for-medical-and-precision-equipment/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 61196-1 / ISO 10993 / UL Style 1354. Values for your order are confirmed in the quotation.
Group | Parameter | Value
Key data | Impedance Z₀ | 50 Ω or 75 Ω ± (Custom 60–90 Ω probe matching)
Key data | Conductor gauge | 36 to 46 AWG (7-strand high-flex Ag-Cu alloy)
Key data | Single coax OD | 0.22 to 0.55 mm (Ultra-thin PFA jacket (0.025 mm wall))
Key data | Capacitance | 40 to 65 pF/m (Foamed PFA (78 % VoP) or solid PFA)
Key data | Flex endurance | > 1,000,000 (± 90° probe strain-relief flex test)
Key data | Bundle counts | 16 to 256 chann (Biocompatible medical PVC/TPU sheath)
Electrical | Characteristic impedance Z₀ | 50 Ω ± 3 Ω or 75 Ω ± 3 Ω (custom 60–85 Ω)
Electrical | Nominal capacitance (50 Ω / 75 Ω) | 95 pF/m (50 Ω) / 40–65 pF/m (75 Ω Foam/PFA)
Electrical | Velocity of propagation (VoP) | 70 % (Solid PFA) / 78–82 % (Foamed PFA)
Electrical | Conductor DC resistance (40 / 42 | ≤ 3.6 / ≤ 5.8 / ≤ 9.2 Ω/m at 20 °C
Electrical | Max. operating pulse voltage | 100–200 V peak-to-peak (piezo transmit)
Electrical | Dielectric withstand test | 500 V rms at 60 Hz
Thermal | Single micro-coax temperature range | -40 °C to +200 °C (PFA fluoropolymer)
Thermal | Medical bundle outer sheath range | -20 °C to +80 °C (medical PVC/TPU) / +150
Thermal | Hot-bar & laser termination heat | Withstands +260 °C soldering without
Thermal | Sterilization compatibility | EtO, STERRAD H₂O₂ plasma & wipe disinfectants
Thermal | Autoclave steam option | +134 °C steam autoclave (silicone outer
Thermal | Flammability rating | UL 1354 / UL VW-1
Mechanical | Conductor gauge range | 36, 38, 40, 42, 44, and 46 AWG (7-strand)
Mechanical | Single micro-coax outer diameter | 0.22 mm (46 AWG) to 0.55 mm (36 AWG)
Mechanical | Complete 128-ch bundle diameter | 7.2 mm to 9.2 mm (ultra-limp medical trunk)
Mechanical | Min. bend radius (single / bundle) | 4 × D (single coax) / 6 × D (probe trunk)
Mechanical | Flex fatigue life (± 90° bend test) | > 1,000,000 cycles (200 g load)
Mechanical | Mass termination pitch | Compatible with 0.25–0.50 mm hot-bar PCBs
Resistance | Biocompatibility (outer sheath) | ISO 10993-5 (cytotoxicity) & -10 (irritation)
Resistance | Hospital disinfectant resistance | Resistant to IPA, Cidex, bleach &
Resistance | Acoustic / triboelectric noise | Low-noise lubricated PTFE binder & shield
Resistance | Inter-channel RF crosstalk | < -65 dB between adjacent bundle channels
Resistance | Tensile break strength (alloy core) | > 650 MPa high-strength Cu-Sn/Ag alloy
Resistance | RoHS / REACH / MDR status | RoHS 3 & EU MDR phthalate-free
Standards | Coaxial transmission standard | IEC 61196-1
Standards | UL appliance wiring style | UL Style 1354 / 1577 / 1911 (30 V, +200 °C)
Standards | Biocompatibility standard | ISO 10993-1 / -5 / -10 (USP Class VI)
Standards | Medical electrical safety | IEC 60601-1 patient-applied part isolation
Standards | Conductor standard | ASTM B298 silver-plated copper alloy
Standards | Quality certification | EN 10204 3.1 + 100 % TDR & capacitance report
Identification | Single micro-coax PFA colours | 10 EIA colours (translucent/solid PFA)
Identification | Sub-bundle group binders | Colour-coded ePTFE threads (8/16-ch groups)
Identification | Medical trunk outer sheath | Medical grey (RAL 7035/7047), white, or black
Identification | Surface legend | Unprinted hospital wipe-clean or custom OEM
Identification | Delivery format | Precision bobbins (single coax) or 16–256
Identification | Batch traceability | Medical device lot traceability

## Construction
Micro-Coaxial Cable for Medical and Precision Equipment – Construction
TEBAOFLEX ¦ https://tebaoflex.com/cables/micro-coaxial-cable-for-medical-and-precision-equipment/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 61196-1 / ISO 10993 / UL Style 1354. Values for your order are confirmed in the quotation.
No. | Layer | Material | Reference
1 | Ultra-fine 7-strand Ag-Cu alloy | 7 × 0.016 mm (46 AWG) to 7 × 0.050 mm (36 AWG) silver-plated high-tensile Cu-Sn/Cu-Ag | ASTM B298 / Custom
2 | Thin-wall solid or cellular foamed PFA | Precision micro-extruded perfluoroalkoxy (PFA, 70 % VoP) or low-capacitance foamed PFA | UL 1354 / IEC 61196
3 | Micro-served spiral or braided shield | 0.025–0.030 mm (50–48 AWG) silver-plated copper alloy spiral serve (≥ 92 % coverage) for | IEC 61196-1
4 | Laser-strippable PFA micro-jacket (+ | 0.020–0.030 mm colour-coded PFA jacket on each coax; 16–256 coax assemblies over-braided | ISO 10993-5/-10
| 40 AWG 50 Ω / 75 Ω Micro-Coax (OD 0.32–0.38 mm) | Workhorse gauge for 64, 128, and 192-channel cart-based diagnostic ultrasound transducer assemblies.
| 42–44 AWG Low-Cap 75 Ω Foamed PFA (≤ 45 pF/m) | Minimizes capacitive loading on piezoelectric crystal elements to maximize transducer sensitivity and
| 46 AWG Sub-Miniature Catheter Coax (OD 0.22 mm) | Ultra-fine micro-coax for intracardiac echocardiography (ICE), IVUS catheters, and flexible video

## Quality and supply
Micro-Coaxial Cable for Medical and Precision Equipment – Testing, approvals and supply
TEBAOFLEX ¦ https://tebaoflex.com/cables/micro-coaxial-cable-for-medical-and-precision-equipment/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 61196-1 / ISO 10993 / UL Style 1354. Values for your order are confirmed in the quotation.
Item | Detail | Frequency / scope
Routine test | 100 % capacitance, TDR impedance & spark test — 50/75 Ω ± 3 Ω; capacitance within ± 2 pF/m across all channels | Every bobbin / bundle
Routine test | 500 V AC dielectric & channel-to-channel shorts — Zero breakdown or open/short across 100 % of bundled cores | Every bundle
Sample test | Laser-strip cleanliness & hot-bar solderability — Clean UV/CO₂ laser window stripping without shield nicking | Per production lot
Type test | 1,000,000-cycle ± 90° probe flex & ISO 10993 — < 5 % resistance shift after 1M flexes; passes biocompatibility | Qualification
Standard / approval | UL Style 1354 / IEC 61196
Standard / approval | ISO 10993-5 / -10
Standard / approval | RoHS 3 / EU MDR
Company certificate | VDE
Company certificate | UL
Company certificate | cUL
Company certificate | CE
Company certificate | KTL
Company certificate | MSHA
Company certificate | MA
Company certificate | ISO 9001
Company certificate | ISO 14001
Company certificate | ISO 45001
Sheath marking | TEBAOFLEX MEDICAL MICRO-COAXIAL [36-46 AWG]
Supply | Delivery form: Spools, reels or export wooden drums
Supply | Standard reel lengths: 100 m · 305 m (1000 ft) · 500 m · 1000 m
Supply | MOQ: 100 m (stock) · 305–500 m (production)
Supply | Lead time: 1–3 weeks ex-works
Supply | Incoterms: EXW, FCA, FOB, CIF, DAP