## Request a quote
Request a quotation for Mineral-Insulated Thermocouple Cable
Application engineer · MIMS pyrometry & furnace sensors
Enter the length you need next to each size in the Sizes sheet (column A) and send this file to:
Email: info@tebaoflex.com
WhatsApp / phone: +86 185 9591 3361
Online quote form: https://tebaoflex.com/cables/mineral-insulated-thermocouple-cable/#template-selection
An application engineer replies within 2 business hours.

## Sizes
Mineral-Insulated Thermocouple Cable – Dimensions, wall thickness & weights (16)
TEBAOFLEX ¦ https://tebaoflex.com/cables/mineral-insulated-thermocouple-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 61515 / ASTM E585 / AMS 2750. Values for your order are confirmed in the quotation.
Enter the length you need next to each size and send this file to info@tebaoflex.com or WhatsApp +86 185 9591 3361.
Your length (m) | Order designation | Type & poles | Sheath alloy | Sheath OD mm | Wall nom. mm | Wire Ø mm ≈ | Weight ≈ kg/km | Max. sheath temp. | Series
| MI-K-1P-1.0-A600 | Type K · Simplex (2-core) | Alloy 600 (NiCr15Fe) | 1 | 0.12 | 0.17 | 6.2 | +1100 °C | Type K (NiCr / NiAl) Class 1 — Alloy 600 & AISI 310 Sheath
| MI-K-1P-1.5-A600 | Type K · Simplex (2-core) | Alloy 600 (NiCr15Fe) | 1.5 | 0.18 | 0.26 | 13.8 | +1100 °C | Type K (NiCr / NiAl) Class 1 — Alloy 600 & AISI 310 Sheath
| MI-K-1P-3.0-A600 | Type K · Simplex (2-core) | Alloy 600 (NiCr15Fe) | 3 | 0.36 | 0.51 | 55 | +1150 °C | Type K (NiCr / NiAl) Class 1 — Alloy 600 & AISI 310 Sheath
| MI-K-1P-4.5-A600 | Type K · Simplex (2-core) | Alloy 600 (NiCr15Fe) | 4.5 | 0.54 | 0.77 | 124 | +1150 °C | Type K (NiCr / NiAl) Class 1 — Alloy 600 & AISI 310 Sheath
| MI-K-1P-6.0-A600 | Type K · Simplex (2-core) | Alloy 600 (NiCr15Fe) | 6 | 0.72 | 1.02 | 220 | +1150 °C | Type K (NiCr / NiAl) Class 1 — Alloy 600 & AISI 310 Sheath
| MI-K-2P-3.0-A600 | Type K · Duplex (4-core) | Alloy 600 (NiCr15Fe) | 3 | 0.36 | 0.39 | 57 | +1150 °C | Type K (NiCr / NiAl) Class 1 — Alloy 600 & AISI 310 Sheath
| MI-K-2P-6.0-A600 | Type K · Duplex (4-core) | Alloy 600 (NiCr15Fe) | 6 | 0.72 | 0.78 | 228 | +1150 °C | Type K (NiCr / NiAl) Class 1 — Alloy 600 & AISI 310 Sheath
| MI-K-1P-6.0-SS310 | Type K · Simplex (2-core) | AISI 310 (25Cr-20Ni) | 6 | 0.72 | 1.02 | 216 | +1100 °C | Type K (NiCr / NiAl) Class 1 — Alloy 600 & AISI 310 Sheath
| MI-N-1P-3.0-A600 | Type N · Simplex (2-core) | Alloy 600 (NiCr15Fe) | 3 | 0.36 | 0.51 | 55 | +1150 °C | Type N (NiCrSi / NiSi) Class 1 — Alloy 600 & NiCrSi Sheath (to +1250 °C)
| MI-N-1P-4.5-A600 | Type N · Simplex (2-core) | Alloy 600 (NiCr15Fe) | 4.5 | 0.54 | 0.77 | 124 | +1200 °C | Type N (NiCrSi / NiSi) Class 1 — Alloy 600 & NiCrSi Sheath (to +1250 °C)
| MI-N-1P-6.0-NiCrSi | Type N · Simplex (2-core) | NiCrSi-type sheath | 6 | 0.72 | 1.02 | 222 | +1250 °C | Type N (NiCrSi / NiSi) Class 1 — Alloy 600 & NiCrSi Sheath (to +1250 °C)
| MI-N-2P-6.0-NiCrSi | Type N · Duplex (4-core) | NiCrSi-type sheath | 6 | 0.72 | 0.78 | 230 | +1250 °C | Type N (NiCrSi / NiSi) Class 1 — Alloy 600 & NiCrSi Sheath (to +1250 °C)
| MI-J-1P-3.0-SS316L | Type J · Simplex (2-core) | AISI 316L Stainless | 3 | 0.36 | 0.51 | 54 | +750 °C | Type J, T & E Class 1 — AISI 316L Stainless Steel Sheath
| MI-J-1P-6.0-SS316L | Type J · Simplex (2-core) | AISI 316L Stainless | 6 | 0.72 | 1.02 | 215 | +750 °C | Type J, T & E Class 1 — AISI 316L Stainless Steel Sheath
| MI-T-1P-3.0-SS316L | Type T · Simplex (2-core) | AISI 316L Stainless | 3 | 0.36 | 0.51 | 54 | +400 °C | Type J, T & E Class 1 — AISI 316L Stainless Steel Sheath
| MI-E-1P-3.0-SS316L | Type E · Simplex (2-core) | AISI 316L Stainless | 3 | 0.36 | 0.51 | 54 | +850 °C | Type J, T & E Class 1 — AISI 316L Stainless Steel Sheath
Custom diameters from 0.25 mm to 10.8 mm and heavy-wall (15 % OD) petrochemical grades are manufactured to order.

## IEC 60584-1 Class 1 tolerances
Mineral-Insulated Thermocouple Cable – IEC 60584-1 Class 1 tolerances & sheath selection guide (5)
TEBAOFLEX ¦ https://tebaoflex.com/cables/mineral-insulated-thermocouple-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 61515 / ASTM E585 / AMS 2750. Values for your order are confirmed in the quotation.
Thermocouple type | IEC Class 1 tolerance | Recommended sheath alloy | Optimal furnace atmosphere
Type K (NiCr / NiAl) | ±1.5 °C (−40…+375 °C) or ±0.004·¦t¦ (+375…+1000 °C) | Alloy 600 or AISI 310 | Oxidising or inert (avoid sulphur > +550 °C on Alloy 600)
Type N (NiCrSi / NiSi) | ±1.5 °C (−40…+375 °C) or ±0.004·¦t¦ (+375…+1000 °C) | NiCrSi sheath or Alloy 600 | High-stability oxidising, vacuum & AMS 2750 pyrometry to +1250 °C
Type J (Fe / CuNi) | ±1.5 °C (−40…+375 °C) or ±0.004·¦t¦ (+375…+750 °C) | AISI 316L or AISI 321 | Reducing, vacuum, plastics extrusion & inert to +750 °C
Type T (Cu / CuNi) | ±0.5 °C (−40…+125 °C) or ±0.004·¦t¦ (+125…+350 °C) | AISI 316L | Cryogenic (−200 °C), pharmaceutical autoclaves & food to +400 °C
Type E (NiCr / CuNi) | ±1.5 °C (−40…+375 °C) or ±0.004·¦t¦ (+375…+800 °C) | AISI 316L or Alloy 600 | High-EMF cryogenic, power generation & oxidising to +850 °C
All MIMS coils are delivered hermetically sealed at both ends to prevent hygroscopic MgO moisture absorption during storage.

## Technical data
Mineral-Insulated Thermocouple Cable – Technical data
TEBAOFLEX ¦ https://tebaoflex.com/cables/mineral-insulated-thermocouple-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 61515 / ASTM E585 / AMS 2750. Values for your order are confirmed in the quotation.
Group | Parameter | Value
Key data | Calibrations | K, N, J, T, E (Simplex (1P), Duplex (2P) & 3P)
Key data | Max. temp. | Up to +1250 °C (Alloy 600 / AISI 310 / NiCrSi)
Key data | Insulation | Compacted MgO (≥ 99.4 % purity (or nuclear 99.7%))
Key data | Sheath Ø (OD) | 0.5 to 8.0 mm (wall thickness ≥ 10–12 % of OD)
Key data | Bend radius | 2 × OD (min.) (formable without losing IR)
Key data | Pressure rating | Up to 400 MPa (hermetic seamless drawn tube)
Electrical | Thermocouple calibrations | Type K, N, J, T, E (IEC 60584-1 Class 1 standard; Class 2 optional)
Electrical | Conductor configurations | Simplex (1P / 2-wire), Duplex (2P / 4-wire), Triplex (3P / 6-wire)
Electrical | Insulation resistance @ 20 °C | ≥ 1,000 MΩ·m (1 GΩ·m) at 500 V DC (≥ 100 V DC for OD < 1.5 mm)
Electrical | Insulation resistance @ +1000 °C | ≥ 0.5 MΩ·m (Type K / N high-purity MgO ≥ 99.4 %)
Electrical | Dielectric withstand voltage | 500 V rms AC (1.0–1.5 mm OD) · 1,000 V rms AC (≥ 2.0 mm OD) for 1 min
Thermal | Alloy 600 (NiCr15Fe / UNS N06600) | Max. continuous +1150 °C (oxidising/inert/vacuum; avoid S > +550 °C)
Thermal | NiCrSi-type high-temp sheath | Max. continuous +1250 °C (matched thermal expansion & chemistry for Type N)
Thermal | AISI 310 (25Cr-20Ni / 1.4841) | Max. continuous +1100 °C (resists oxidising & sulphurous combustion gases)
Thermal | AISI 316L (17Cr-12Ni-2Mo / 1.4404) | Max. continuous +850 °C (chemical, pharma, marine & food service)
Mechanical | Standard outer diameters (OD) | 0.50, 1.0, 1.5, 2.0, 3.0, 4.5, 6.0, 8.0 mm (and 1/16" to 1/4")
Mechanical | Nominal sheath wall thickness | 10 % to 12 % of cable OD (minimum ≥ 8 % of OD)
Mechanical | Nominal conductor diameter | Simplex: ≈ 17 % of OD · Duplex: ≈ 13 % of OD
Mechanical | Minimum bend radius | 2 × OD around mandrel without sheath cracking or IR loss
Standards | Primary MIMS cable standards | IEC 61515, ASTM E585 / E585M & ASTM E608
Standards | Pyrometry & calibration standards | IEC 60584-1 Class 1, ANSI/ISA MC96.1 Special Limits & AMS 2750
Identification | Coil / spool delivery | Bright-annealed moisture-sealed coils with brazed/resin-sealed ends
Identification | Traceability | Heat/lot number, EMF calibration certificate at test points, and PMI report
Identification | Customs tariff (HS) | 8544.49 / 9033.00

## Construction
Mineral-Insulated Thermocouple Cable – Construction
TEBAOFLEX ¦ https://tebaoflex.com/cables/mineral-insulated-thermocouple-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 61515 / ASTM E585 / AMS 2750. Values for your order are confirmed in the quotation.
No. | Layer | Material | Reference
1 | Solid Class 1 thermocouple alloy conductors | Simplex (2-wire), Duplex (4-wire) or Triplex (6-wire) Type K (NiCr/NiAl), Type N (NiCrSi/NiSi), Type J (Fe/CuNi), Type T (Cu/CuNi) or Type E (NiCr/CuNi)
2 | High-density compacted Magnesium Oxide (MgO) | Refractory MgO dielectric powder (≥ 99.4 % purity, low Fe₂O₃/B₂O₃/SiO₂ impurity content) compacted to > 75 % theoretical density
3 | Seamless cold-drawn & bright-annealed metallic sheath | Alloy 600 (NiCr15Fe, UNS N06600), NiCrSi sheath alloy, AISI 310 (25Cr-20Ni), AISI 316L (UNS S31603) or AISI 321 (nominal wall 10–12 % of OD)
| Alloy 600 Sheath (UNS N06600 · to +1150 °C) | 72Ni-15Cr-8Fe superalloy resisting oxidation, chloride SCC and carburisation ( avoid sulphur > +550 °C)
| NiCrSi Sheath for Type N (to +1250 °C) | Matched Nicrosil-type Cr-Si-Ni sheath eliminating Ni/Cr vapour cross-contamination of Type N conductors
| AISI 310 / AISI 316L Stainless Sheath | AISI 310 (25Cr-20Ni) for sulphurous furnace gases to +1100 °C; AISI 316L for chemical & food process to +850 °C

## Quality and supply
Mineral-Insulated Thermocouple Cable – Testing, approvals and supply
TEBAOFLEX ¦ https://tebaoflex.com/cables/mineral-insulated-thermocouple-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on IEC 61515 / ASTM E585 / AMS 2750. Values for your order are confirmed in the quotation.
Item | Detail | Frequency / scope
Routine test | Insulation resistance (IR) & AC dielectric test (IEC 61515) — IR ≥ 1,000 MΩ·m at 20 °C (500 V DC); AC dielectric test without breakdown | Every manufactured coil
Routine test | Multi-point thermoelectric EMF calibration & sheath ductility test — IEC 60584-1 Class 1 EMF tolerance; 2 × OD mandrel bend & helium/dye-penetrant sheath check | Every production lot
Standard / approval | IEC 61515 & ASTM E585 Certified
Standard / approval | IEC 60584-1 Class 1 & AMS 2750 Compliant
Standard / approval | EN 10204 3.1 Material & Calibration Cert.
Company certificate | VDE
Company certificate | UL
Company certificate | cUL
Company certificate | CE
Company certificate | KTL
Company certificate | MSHA
Company certificate | MA
Company certificate | ISO 9001
Company certificate | ISO 14001
Company certificate | ISO 45001
Sheath marking | Tagged Stainless Spool / Coil: TEBAOFLEX MI-<type>-<OD>-<sheath> IEC 61515 CLASS 1 LOT#
Supply | Delivery form: Non-returnable fumigated wooden drums
Supply | Standard lengths: 500 m · 1000 m standard drums (cut lengths on request)
Supply | MOQ: 500 m (stock sizes) · 1000 m (custom)
Supply | Lead time: 2–4 weeks ex-works
Supply | Incoterms: EXW, FCA, FOB, CIF, DAP