## Request a quote
Request a quotation for SABmed S Connecting Cable
Medical diagnostic & surgical handpiece cable engineer
Enter the length you need next to each size in the Sizes sheet (column A) and send this file to:
Email: info@tebaoflex.com
WhatsApp / phone: +86 185 9591 3361
Online quote form: https://tebaoflex.com/cables/sabmed-s-biocompatible-silicone-medical-device-diagnostic-connecting-cable/#template-selection
An application engineer replies within 2 business hours.

## Sizes
SABmed S Connecting Cable – Cores × mm², EMC screen, core insulation, OD & weight (12)
TEBAOFLEX ¦ https://tebaoflex.com/cables/sabmed-s-biocompatible-silicone-medical-device-diagnostic-connecting-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on ISO 10993-1 / -5 / -10 · USP Class VI · IEC 60601-1 · DIN EN ISO 17665-1. Values for your order are confirmed in the quotation.
Enter the length you need next to each size and send this file to info@tebaoflex.com or WhatsApp +86 185 9591 3361.
Your length (m) | Order designation | Cores × Cross-Section | EMC Screen & Core Wall | Rated U V AC | Nom. OD mm | Cu Index kg/km | Weight kg/km | Series
| SABMED-S-2X0.25 | 2 × 0.25 mm² (24 AWG) | Unscreened · SiR Cores | 300 | 3.8 | 4.8 | 18.5 | SABmed S Unscreened Multi-Core Medical Connecting Cables (ISO 10993 / USP Class VI)
| SABMED-S-3X0.50 | 3 × 0.50 mm² (20 AWG) | Unscreened · SiR Cores | 500 | 4.9 | 14.4 | 34 | SABmed S Unscreened Multi-Core Medical Connecting Cables (ISO 10993 / USP Class VI)
| SABMED-S-4X0.25 | 4 × 0.25 mm² (24 AWG) | Unscreened · SiR Cores | 300 | 4.3 | 9.6 | 26.5 | SABmed S Unscreened Multi-Core Medical Connecting Cables (ISO 10993 / USP Class VI)
| SABMED-S-6X0.25 | 6 × 0.25 mm² (24 AWG) | Unscreened · FEP Cores | 300 | 4.6 | 14.4 | 31 | SABmed S Unscreened Multi-Core Medical Connecting Cables (ISO 10993 / USP Class VI)
| SABMED-S-3X1.00 | 3 × 1.00 mm² (18 AWG) | Unscreened · SiR Cores | 500 | 6.2 | 28.8 | 58 | SABmed S Unscreened Multi-Core Medical Connecting Cables (ISO 10993 / USP Class VI)
| SABMED-SC-4X0.14 | 4 × 0.14 mm² (26 AWG) | ≥ 85% SnCu Braid · FEP | 300 | 3.9 | 11.2 | 23.5 | SABmed S-C EMC-Screened & Hybrid Surgical Motor / Patient Diagnostic Cables
| SABMED-SC-6X0.14 | 6 × 0.14 mm² (26 AWG) | ≥ 85% SnCu Braid · FEP | 300 | 4.4 | 14.5 | 29.8 | SABmed S-C EMC-Screened & Hybrid Surgical Motor / Patient Diagnostic Cables
| SABMED-SC-8X0.14 | 8 × 0.14 mm² (26 AWG) | ≥ 85% SnCu Braid · FEP | 300 | 4.8 | 17.8 | 35.5 | SABmed S-C EMC-Screened & Hybrid Surgical Motor / Patient Diagnostic Cables
| SABMED-SC-4X0.25 | 4 × 0.25 mm² (24 AWG) | ≥ 85% SnCu Braid · FEP | 300 | 4.4 | 16 | 31.2 | SABmed S-C EMC-Screened & Hybrid Surgical Motor / Patient Diagnostic Cables
| SABMED-SC-12X0.14 | 12 × 0.14 mm² (26 AWG) | ≥ 85% SnCu Braid · FEP | 300 | 5.5 | 24.5 | 46.8 | SABmed S-C EMC-Screened & Hybrid Surgical Motor / Patient Diagnostic Cables
| SABMED-SC-3X0.50+6X0.14 | 3×0.50 + 6×0.14 mm² Motor | ≥ 85% SnCu Braid · FEP | 300 | 5.8 | 31 | 54 | SABmed S-C EMC-Screened & Hybrid Surgical Motor / Patient Diagnostic Cables
| SABMED-SC-3X0.75+8X0.14 | 3×0.75 + 8×0.14 mm² Handpiece | ≥ 85% SnCu Braid · FEP | 300 | 6.4 | 42.5 | 68.5 | SABmed S-C EMC-Screened & Hybrid Surgical Motor / Patient Diagnostic Cables
Compatible with medical push-pull circular connectors and silicone liquid-injection-moulded (LIM) bend reliefs.

## Technical data
SABmed S Connecting Cable – Technical data
TEBAOFLEX ¦ https://tebaoflex.com/cables/sabmed-s-biocompatible-silicone-medical-device-diagnostic-connecting-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on ISO 10993-1 / -5 / -10 · USP Class VI · IEC 60601-1 · DIN EN ISO 17665-1. Values for your order are confirmed in the quotation.
Group | Parameter | Value
Key data | Biocompatibility | ISO 10993/USP VI (Skin-contact non-cytotoxic SiR)
Key data | Core count | 2 to 24+ Cores (0.14 mm² (26 AWG) to 2.5 mm²)
Key data | EMC option | Unshld or ≥85% (Tinned or silver-plated Cu braid)
Key data | Sterilization | +134°C Autoclave (Steam, EtO, H2O2 plasma & wipes)
Key data | Voltage rating | 300 / 500 V AC (1.5 kV – 2.5 kV AC dielectric test)
Key data | Temperature | −50 to +180 °C (Limp, non-stick matte surface)
Electrical | Nominal voltage (U0/U) | 300 / 300 V (≤ 0.25 mm²) · 300 / 500 V (≥ 0.34 mm²)
Electrical | Routine AC test voltage | 1,500 V AC to 2,500 V AC core-to-core & core-to-screen
Electrical | Insulation resistance | ≥ 1,000 MΩ·km (silicone cores) · ≥ 5,000 MΩ·km (FEP cores)
Thermal | Operating temperature | −50 °C to +180 °C (peaks to +200 °C)
Thermal | Steam autoclave sterilization | +134 °C to +138 °C at 2.1 bar (ISO 17665-1 reusable instrument grade)
Mechanical | Core cross-sections | 0.08, 0.14, 0.25, 0.34, 0.50, 0.75, 1.0, 1.5 & 2.5 mm² (Class 6)
Mechanical | Min. bend radius | 4 × OD static · 7.5 × OD dynamic clinical flexing
Resistance | Biocompatibility | ISO 10993-1, ISO 10993-5 (Cytotoxicity), ISO 10993-10 · USP Class VI
Resistance | Disinfectant resistance | Isopropanol, ethanol, aldehyde-free & quaternary ammonium hospital wipes
Resistance | Halogen & plasticizer free | 100 % platinum-cured silicone · zero DEHP/phthalates, latex or halogens
Standards | Medical safety & biocompatibility | IEC 60601-1 · ISO 10993-1/-5/-10 · USP Class VI · EU MDR 2017/745 material support
Standards | Environmental compliance | RoHS 2011/65/EU, (EU) 2015/863 & REACH compliant
Identification | Core identification | DIN 47100 colour code
Identification | Outer sheath colours | Light Grey (RAL 7035), Pure White (RAL 9010), Surgical Blue, Black, Mint Green
Identification | Customs tariff (HS) | 8544.49

## Construction
SABmed S Connecting Cable – Construction
TEBAOFLEX ¦ https://tebaoflex.com/cables/sabmed-s-biocompatible-silicone-medical-device-diagnostic-connecting-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on ISO 10993-1 / -5 / -10 · USP Class VI · IEC 60601-1 · DIN EN ISO 17665-1. Values for your order are confirmed in the quotation.
No. | Layer | Material | Reference
1 | Extra-fine Class 6 tinned or silver-plated copper (Ø 0.05 mm) | Finely stranded 0.05 mm tinned copper (or silver-plated copper / Cu-alloy on 28–26 AWG micro-coax/data pairs)
2 | Thin-wall FEP / ETFE fluoropolymer or platinum silicone cores | DIN 47100 colour-coded FEP/ETFE (for compact OD and high steam resistance) or platinum-cured silicone primary core insulation
3 | PTFE / non-woven medical separating slip tape | Helical fluoropolymer tape over the cabled core bundle
4 | ≥ 85 % tinned copper braided EMC screen (SABmed S-C variants) | High-flexibility fine-wire tinned copper braid (plus optional aluminised foil on twisted-pair data builds)
5 | Platinum-cured ISO 10993 / USP Class VI medical silicone jacket | Addition-crosslinked biocompatible silicone rubber with velvet-matte anti-tack surface finish in RAL 7035 Grey, White, Blue or Black
| SABmed S Unscreened (2 to 18 Cores, 0.14–2.5 mm²) | Ultra-limp unshielded multi-core medical cable for surgical foot switches, therapy applicators, power pods and general sensor leads
| SABmed S-C Screened & Hybrid (2 to 24 Cores + 85 % Cu Braid) | EMC-braided multi-core, twisted-pair or hybrid power+signal (e.g. 3×0.75 + 6×0.14 mm²) cable for surgical handpiece motors & patient monitors

## Quality and supply
SABmed S Connecting Cable – Testing, approvals and supply
TEBAOFLEX ¦ https://tebaoflex.com/cables/sabmed-s-biocompatible-silicone-medical-device-diagnostic-connecting-cable/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on ISO 10993-1 / -5 / -10 · USP Class VI · IEC 60601-1 · DIN EN ISO 17665-1. Values for your order are confirmed in the quotation.
Item | Detail | Frequency / scope
Routine test | 100 % spark test & 1.5–2.5 kV AC dielectric withstand test — Verify core-to-core and core-to-screen insulation integrity across every manufactured length | 100 % of length
Sample test | ISO 10993-5 cytotoxicity, -10 skin irritation & USP Class VI certification — Full biological evaluation certificate for patient-contact platinum-cured outer sheath | Compound qualification
Sample test | +134 °C steam autoclave & clinical flex-life endurance test — Verify mechanical elasticity, colour stability and electrical continuity after sterilization cycles | Production audit
Standard / approval | ISO 10993-1 / -5 / -10 & USP Class VI Biocompatible
Standard / approval | IEC 60601-1 & ISO 17665-1 (+134 °C Autoclave) Ready
Company certificate | VDE
Company certificate | UL
Company certificate | cUL
Company certificate | CE
Company certificate | KTL
Company certificate | MSHA
Company certificate | MA
Company certificate | ISO 9001
Company certificate | ISO 14001
Company certificate | ISO 45001
Sheath marking | Smooth Wipe-Clean Sterilizable Surface (Unprinted or Biocompatible Laser Legend)
Supply | Delivery form: Cleanroom double-bagged spools or overmould-ready cut assemblies
Supply | Regulatory dossier: ISO 10993-5/-10 test summary, USP Class VI cert, RoHS/REACH & MDR declaration
Supply | Lead time: 2–4 weeks ex-works