## Request a quote
Request a quotation for WDT / WQT Cryogenic Twisted Pair
Application engineer · instrumentation & sensing cables
Enter the length you need next to each size in the Sizes sheet (column A) and send this file to:
Email: info@tebaoflex.com
WhatsApp / phone: +86 185 9591 3361
Online quote form: https://tebaoflex.com/cables/wdt-wqt-phosphor-bronze-and-manganin-loom-woven-cryogenic-twisted-pair-wire/#template-selection
An application engineer replies within 2 business hours.

## Sizes
WDT / WQT Cryogenic Twisted Pair – Cryogenic electrical & thermal properties (4)
TEBAOFLEX ¦ https://tebaoflex.com/cables/wdt-wqt-phosphor-bronze-and-manganin-loom-woven-cryogenic-twisted-pair-wire/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on NEMA MW 1000 (MW 16-C / MW 28-C) / ASTM B103. Values for your order are confirmed in the quotation.
Enter the length you need next to each size and send this file to info@tebaoflex.com or WhatsApp +86 185 9591 3361.
Your length (m) | Order designation | Configuration & alloy | Bundle Ø mm | R @ 295K Ω/m | R @ 4.2K Ω/m | k @ 4.2K W/m·K | Series
| WDT-32 (Phos-Bronze Pair) | 2-Wire Pair × 32 AWG Phosphor-Bronze | 0.47 | 19 | 16.5 | 2.2 | Cryogenic Twisted Pair (WDT) & 4-Wire Quad (WQT) Sensor Leads
| WDT-36 (Phos-Bronze Pair) | 2-Wire Pair × 36 AWG Phosphor-Bronze | 0.31 | 48 | 41.8 | 2.2 | Cryogenic Twisted Pair (WDT) & 4-Wire Quad (WQT) Sensor Leads
| WQT-32 (Phos-Bronze Quad) | 4-Wire Quad (2P) × 32 AWG Phosphor-Bronze | 0.94 | 19 | 16.5 | 2.2 | Cryogenic Twisted Pair (WDT) & 4-Wire Quad (WQT) Sensor Leads
| WQT-36 (Phos-Bronze Quad) | 4-Wire Quad (2P) × 36 AWG Phosphor-Bronze | 0.62 | 48 | 41.8 | 2.2 | Cryogenic Twisted Pair (WDT) & 4-Wire Quad (WQT) Sensor Leads
Resistance values are per single conductor. Manganin versions (WDT-MN-36 / WQT-MN-36) are also stocked.

## Technical data
WDT / WQT Cryogenic Twisted Pair – Technical data
TEBAOFLEX ¦ https://tebaoflex.com/cables/wdt-wqt-phosphor-bronze-and-manganin-loom-woven-cryogenic-twisted-pair-wire/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on NEMA MW 1000 (MW 16-C / MW 28-C) / ASTM B103. Values for your order are confirmed in the quotation.
Group | Parameter | Value
Key data | Configurations | WDT / WQT (Twisted pair or 4-wire quad)
Key data | Alloys | PhBr / Manganin (CuSn5 (C51000) or Cu86Mn12Ni2)
Key data | Wire gauges | 32 or 36 AWG (0.202 mm or 0.127 mm bare Ø)
Key data | Temp. range | 10 mK to 400 K (-273.14 °C to +125 °C)
Key data | Thermal cond. | 2.2 W/m·K @ 4 K (100× lower heat load than Cu)
Key data | Non-magnetic | Zero Fe/Ni (PhBr (Safe in 20+ Tesla magnets)
Electrical | DC resistance @ 295 K (PhBr) | 19.0 Ω/m (32 AWG) · 48.0 Ω/m (36 AWG) per single conductor
Electrical | DC resistance @ 4.2 K (PhBr) | 16.5 Ω/m (32 AWG) · 41.8 Ω/m (36 AWG) — RRR ≈ 1.15
Electrical | DC resistance @ 295 K (Mang.) | 13.7 Ω/m (32 AWG) · 34.7 Ω/m (36 AWG) — ΔR < 10 % to 4 K
Electrical | Dielectric breakdown | ≥ 600 V rms (twist test per NEMA MW 1000)
Thermal | Operating temperature | 10 mK (-273.14 °C) to 400 K (+127 °C) · Polyimide to +220 °C
Thermal | Thermal cond. (PhBr) | ≈ 48 W/(m·K) @ 300 K · ≈ 12 W/(m·K) @ 77 K · ≈ 2.2 W/(m·K) @ 4.2 K
Thermal | Thermal cond. (Manganin) | ≈ 22 W/(m·K) @ 300 K · ≈ 6.5 W/(m·K) @ 77 K · ≈ 0.5 W/(m·K) @ 4.2 K
Mechanical | Bare / insulated wire Ø | 32 AWG: 0.202 / 0.236 mm · 36 AWG: 0.127 / 0.155 mm
Mechanical | Finished bundle OD | 0.31–0.47 mm (WDT Pair) · 0.62–0.94 mm (WQT Quad)
Mechanical | Twist pitch | 3.9 to 5.5 twists/cm (10 to 14 twists/inch)
Resistance | Magnetic susceptibility | Phosphor-Bronze: strictly non-ferromagnetic (zero Ni / Fe)
Resistance | UHV outgassing | Polyimide ML passes NASA SP-R-0022A (TML < 1.0 %, CVCM < 0.1 %)
Resistance | Cryo varnish compatibility | Resistant to GE 7031 varnish, Stycast 2850FT & Apiezon N
Standards | Magnet wire standard | NEMA MW 1000 (MW 16-C Polyimide / MW 28-C Polyurethane-Nylon)
Standards | Alloy standard | ASTM B103 (C51000 Phosphor Bronze) & DIN 17471 (Manganin)
Standards | RoHS & REACH | 2011/65/EU & EC 1907/2006 compliant
Identification | WQT 4-wire colour code | Pair 1: Clear + Red (I+/I-) · Pair 2: Green + Blue/Amber (V+/V-)
Identification | Customs tariff (HS) | 8544.1900
Identification | Supply spools | Supplied on 15 m, 30 m, 100 m and 300 m cleanroom spools

## Construction
WDT / WQT Cryogenic Twisted Pair – Construction
TEBAOFLEX ¦ https://tebaoflex.com/cables/wdt-wqt-phosphor-bronze-and-manganin-loom-woven-cryogenic-twisted-pair-wire/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on NEMA MW 1000 (MW 16-C / MW 28-C) / ASTM B103. Values for your order are confirmed in the quotation.
No. | Layer | Material | Reference
1 | Solid 32 AWG or 36 AWG Phosphor-Bronze / Manganin wire | Non-magnetic Phosphor-Bronze CDA 510 (Cu94.8Sn5P0.2) or low-TCR Manganin (Cu86Mn12Ni2) solid wire
2 | Heavy Polyimide (NEMA MW 16-C) or Formvar enamel | Concentric 12–18 µm aromatic Polyimide (ML) or Formvar/Polyurethane-Nylon enamel in Clear/Amber, Red, Green, Blue
3 | Precision-twist duo (WDT) or quad (WQT) lay + bondcoat | Controlled 3.9 to 5.5 twists/cm (10–14 twists/in) cancelling dΦ/dt pickup in swept superconducting magnets
| WDT-32 / WDT-36 (Phosphor-Bronze Twisted Pair) | 2-wire twisted pair (32 or 36 AWG Phosphor-Bronze) for non-magnetic cryo sensors and heaters
| WQT-32 / WQT-36 (Phosphor-Bronze 4-Wire Quad) | 4-wire quad (2 twisted pairs × 32 or 36 AWG Phosphor-Bronze) for 4-terminal Cernox/RuO₂ sensors
| WDT-MN-36 / WQT-MN-36 (Manganin Pair / Quad) | 36 AWG Manganin twisted pair or quad with flat resistance vs temperature (300 K to 4 K)

## Quality and supply
WDT / WQT Cryogenic Twisted Pair – Testing, approvals and supply
TEBAOFLEX ¦ https://tebaoflex.com/cables/wdt-wqt-phosphor-bronze-and-manganin-loom-woven-cryogenic-twisted-pair-wire/ ¦ Data revision 2026-10-20
Typical values for this cable type, based on NEMA MW 1000 (MW 16-C / MW 28-C) / ASTM B103. Values for your order are confirmed in the quotation.
Item | Detail | Frequency / scope
Routine test | 100 % NEMA MW 1000 high-voltage pinhole count & DC resistance check — Zero enamel pinholes per 30 m and R within ±3 % of alloy spec | Every manufactured spool
Type test | 20× liquid nitrogen / liquid helium (77 K / 4.2 K) thermal shock test — Zero enamel crazing and ≥ 600 V rms twisted-pair breakdown retention | Per production lot
Sample test | SQUID magnetometer ferromagnetic impurity screening — Zero ferromagnetic remanence on Phosphor-Bronze (CuSn5) lots | Per alloy melt
Standard / approval | NEMA MW 1000 Compliant
Standard / approval | ASTM B103 C51000 Alloy
Standard / approval | NASA SP-R-0022A Low Outgas
Company certificate | VDE
Company certificate | UL
Company certificate | cUL
Company certificate | CE
Company certificate | KTL
Company certificate | MSHA
Company certificate | MA
Company certificate | ISO 9001
Company certificate | ISO 14001
Company certificate | ISO 45001
Sheath marking | Spool label: TEBAOFLEX WDT / WQT CRYOGENIC PHOSPHOR-BRONZE / MANGANIN
Supply | Delivery form: Fumigated wooden drums, steel reels or precision spools
Supply | Standard lengths: 500 m · 1000 m (or custom cut/spooled lengths)
Supply | MOQ: 100 m (standard) · 500 m (custom)
Supply | Lead time: 2–4 weeks ex-works
Supply | Incoterms: EXW, FOB, CIF, DAP